Appendix No. 1 to the Terms of Reference

Technical parameters and requirements for the delivery of the maskaligner (device for centering and exposing photolithographic patterns)

| No | Parameter | Requirement | Column to be filled in by the tenderer* |
| --- | --- | --- | --- |
| 1. | Type |  | Specify |
| 2. | Manufacturer |  | Specify |
| 3. | Country of origin |  | Specify |
| 4. | Year of production | 2025 | Confirm |
| 5. | Device | Brand new, unused | Confirm |
| 6. | Main application | Centering and exposing photolithographic patterns with line widths > 5 micrometers in layers of photosensitive resist using UV light through masks made of chromium layer on glass or quartz substrate for defining contacts in microelectronic devices and micro-energy storage | Confirm |
| 7. | Functionality | - The device allows for exposing patterns from glass or quartz masks with dimensions of 10 cm x 10 cm, 7.5 cm x 7.5 cm and 5 cm x 5 cm - The device allows for centering patterns from the mask as in pt. 1 to an existing pattern using a table with micrometric shifts | Confirm |
| 8. | Light source | - LED, 365 nm | Confirm |
|  |  | - Exposure time adjustable with a step not longer than 1 s | Specify |
| 9. | Exposure field | Minimum diameter of 100 mm | Specify |
| 10. | Mask holder | - Allowing the use of glass or quartz masks with transverse dimensions of 10 cm x 10 cm and smaller - Using vacuum to immobilize masks | Confirm |
| 11. | Sample table | - Allowing work with samples of dimensions 1 cm x 1 cm and thickness < 0.5 mm as well as samples with a diameter of 2” and 4” - Holding samples using vacuum - With micrometric shifts in X, Y, and R (rotation in plane) axes | Confirm |
| 12. | Optical microscope | - Digital - 3-axis - Minimum with three objectives enabling the observations of the patterns on the mask with effective magnifications not worse than 90x, 300x and 450x. - With an UV-free illuminator - Allowing centering patterns across the entire sample surface | Confirm |
| 13. | Control computer | With display, integrated | Confirm |
| 14. | Equipment provided with the device | Technical equipment necessary to ensure the proper operation and functionality of the mask aligner | Confirm |
| 15. | Acceptance test after installation and commissioning of the device at at the Awarding Entity’s site | The acceptance test includes conducting exposure and development processes of electrical contact patterns using 4” masks provided by the Awarding Entity with critical sized of 5, 10, 15, 20, 25, 30 μm. The acceptable deviation of the width of the exposed pattern in a S1800-class photoresist is 10% | Confirm |
| 16. | Training for minimum 3 persons in the usage of the device at the Awarding Entity’s site | to be assured | Confirm |
| 17. | User manual Polish and English along with software manual and technical documentation in Polish or English | to be assured | Confirm |
| 18. | Spare parts availability minimum 10 years from the date of installation | to be assured | Confirm |
| 19. | Free software updates for a minimum of 5 years from the date of installation | to be assured | Confirm |
| 20. | Warranty and post-warranty service response within 3 days from the notification of a fault | to be assured | Confirm |
| 21. | Technical support for a period of 10 years from the date of installation | to be assured | Confirm |

Appendix No. 2 to the Terms of Reference

Statement of the Contractor

submitted pursuant to Art. 125, item 1 of the Act of September 11 2019

Public Procurement Law (hereinafter: the PPL Act),

THE GROUNDS FOR THE EXCLUSION OF ARTICLE 7 item 1 OF THE ACT

ON SPECIAL SOLUTIONS IN THE SCOPE OF COUNTERACTING THE SUPPORT OF AGGRESSION AGAINST UKRAINE AND SERVING THE PROTECTION OF NATIONAL SECURITY

ON THE CONDITIONS FOR EXCLUSION FROM THE PROCEDURE

For the purposes of the procedure for award of a public contract, entitled …………………………………………………… (name of the procedure), conducted by Łukasiewicz Research Network – Institute of Microelectronics and Photonics (name of the Contracting Authority), I hereby represent as follows:

STATEMENTS ON THE CONTRACTOR:

- I represent that I am not subject to exclusion from the procedure under Art. 108 item 1 and 109 item 1 (4) of the PPL Act and Art. 7 para. 1 (1-3) of the Act of April 13, 2022 on special solutions in the scope of counteracting the support of aggression against Ukraine and serving the protection of national security

or

- that grounds for exclusion from the procedure exist pursuant to Art. …………… of the PPL Act

At the same time, I represent that, in connection with these circumstances, I have taken the following corrective actions pursuant to Art. 110 item 2 of the PPL Act: ………………………………………………………………………………….

STATEMENT CONCERNING THE ENTITY WHOSE RESOURCES WILL BE REFERRED BY THE CONTRACTOR

- I certify that the following entity/s, whose resources I am referring to in these proceedings, i.e. …………………………………………………………………….……………………… (full name of the entity, address, and also depending on the entity: Tax ID/Personal ID, National Court Register/Register of Business Activity Number), is/are not subject of exclusion from the public Contract awarding procedure.

STATEMENT CONCERNING A SUBCONTRACTOR WHICH IS NOT AN ENTITY WHOSE RESOURCES WILL BE REFERRED TO BY THE CONTRACTOR

- I certify that the following entity/s, being the Subcontractor/s: ……………………………………………… (full name of the entity, address, and also depending on the entity: Tax ID/Personal ID, National Court Register/Register of Business Activity Number), is/are not subject of exclusion from the public Contract awarding procedure.

DECLARATION CONCERNING THE INFORMATION PROVIDED:

- I declare that all the information provided in the above statements is up-to-date and truthful and has been presented with full awareness of the consequences of misleading the Contracting Authority when providing the information.

I indicate that the documents confirming these facts, referred to in Chapter VII of the ToR, are available in electronic form at the following internet addresses of public and free (please select):

- https://prod.ceidg.gov.pl

- https://ems.ms.gov.pl

- Other register: ………………………………..

Appendix No. 3 to the Terms of Reference

Contractors who jointly apply for a contract:

………………………………………

………………………………………

(Full name/Company name, address)

Represented by:

………………………………………

(Representative Name, Surname, Position, Title)

STATEMENT OF THE CONTRACTOR

ABOUT THE SCOPE OF PERFORMANCE OF THE CONTRACT BY THE CONTRACTORS

WHO JOINTLY APPLY FOR A CONTRACT

submitted on the basis of art. 117 item 4 of the Act of September 11, 2019 - Public Procurement Law (Official Journal of 2019, item 2019 with later changes)

in the procedure for awarding a public Contract for

………………………………………………………………………………………..

we declare that:

- Contractor …………………………………………………………………………………………………….…………….…..

(name and address name and address of one of the Contractors who jointly apply for a contract, e.g. a consortium member or partner in a civil partnership)

will perform the following services *:…………………………………………

- Contractor ……………………………………………………………………………………………….……………..……..

(name and address name and address of one of the Contractors who jointly apply for a contract, e.g. a consortium member or partner in a civil partnership)

will perform the following services *:…………………………………………

NOTE!

It should be indicated which supplies/services/construction works * included in scope of the subject of the contract are performed by individual contractors (consortium members or partners of a civil partnership)

STATEMENT REGARDING THE INFORMATION PROVIDED

I/We* hereby represent that any and all information provided in this Statement are up-to-date and accurate and that they I/we are fully aware of the consequences of providing the Contracting Authority with misleading information.

*Delete as appropriate

Appendix No. 4 to the Terms of Reference

BID FORM

for: Łukasiewicz Research Network – Institute of Microelectronics and Photonics

| Subject of the Order |  | Delivery of the maskaligner (device for centering and exposing photolithographic patterns) |
| --- | --- | --- |
| 1. | Name and address of the Contractor | ………………………………………………………………………………. ………………………………………………………………………………. |
| 2. | Taxpayer’s Identification Number NIP REGON number: or the relevant foreign identification numbers | ……………………………………………………………………………… …………………………………………………………………………….. ……………………………………………………………………………. |
| 3. | Telephone: E-mail: | ............................................................................... .............................................................................. |
| 4. | Total net price of the bid VAT rate VAT amount Total gross price of the bid | enter: ..................... PLN/EUR/USD/GBP* Say: ................................................................................. ........................................................................................enter: .....................% enter: ..................... PLN/EUR/USD/GBP* enter: ..................... PLN/EUR/USD/GBP* Say: ................................................................................. |
| 5. | Time of performing the Contract: Maximum time: up to 4 months from the date of signing the agreement | enter: ........................................ (months) |
| 6. | Warranty period: at least 12 months | enter:…………………….. (months) |

| 7. | Statement | I hereby represent that the selection of our bid will/will not* result in imposing taxation obligation on the Contracting Authority. The following (name or type of the goods or services) will lead to the emergence of a tax obligation: Value without tax:……………………………………………………. Name of goods:………………………………………………………… |
| --- | --- | --- |
| 8. | Type of Contractor | I declare that I am: - Microentrepreneur - Small enterprise - Medium enterprise - A sole proprietorship - A natural person who does not conduct business activity - Other types: specify ………………………. within the meaning of the Act of March 6, 2018 Entrepreneurs Law. (Note! Mark the "X" in the appropriate box)) |
| 9. | Statement | I hereby represent that I have read the Terms of Reference (together with the Appendices constituting an integral part thereof) and that I accept its content without reservations. |
| 10. | I hereby represent that I have fulfilled the information obligations foreseen in Art. 13 or Art. 14 of the GDPR towards natural persons whose personal data I obtained directly or indirectly in order to apply for awarding a public contract in the present procedure. | confirm:…………………………………….. In the event if the Contractor does not transfer the personal data other than their own or if the information obligation is excluded pursuant to Art. 13, item 4 or Art. 14, item 5 of the GDPR, the Contractor shall not submit this statement (delete the content of the statement e.g. by crossing out). |
| 11. | Persons authorised to represent the Contractor | enter: ........................................... |
| 12. | Parts of the Contract that will be entrusted to subcontractors | Specify the parts of the Contract: ....................... Enter names of companies: ………………………….. |
| 13. | Documents attached to the bid | 1. …………………………………………………………………. 2. …………………………………………………………………. 3. …………………………………………………………………. |

* Delete as appropriate

Appendix No. 6 to the Terms of Reference

STATEMENT

STATEMENT OF THE CONTRACTOR CONFIRMING THE VALID OF INFORMATION PROVIDED 
IN THE STATMENT SPECIFIED IN ART. 125 OF THE PPL ACT, ON THE GROUNDS FOR EXCLUSION

Contractor :…………………………………………………………………………………………………………………………………………….

…………………………………………………………………………………………………….…………………………………………………………

For the purposes of procurement procedure, the subject of which is

…………………………………………………………………………………………………….…………………………………………………………

We declare that the information provided in the statement placed pursuant to Art. 125, item 1 of the PPL Act, with respect to the ground for exclusion from proceeding, stipulated in:

- art. 108 ust. 1 of the PPL Act

- art. 7 para. 1 (1-3) of the Act of April 13, 2022 on special solutions in the scope of counteracting the support of aggression against Ukraine and serving the protection of national security

are up-to-date / are out-of-date*

NOTE: In the previously provided information is not up-to-date, please submit additional information in this regard, in particular specify which data the change relates to and indicate its scope

*delete as appropriate