Appendix A - Description of the subject of the contract for the NAP-XPS system for beamline at the SOLARIS National Synchrotron Radiation Centre

Table of Contents

1.	General information	3

1.1. General specification of the beamline	4

2.	Scope of delivery	6

2.1 Part 1: Near Ambient Pressure XPS system	6

2.1.1. Analysis Chamber for X-ray Photoelectron Spectroscopy	7

2.1.2. Equipment of the Analysis Chamber	9

2.1.3. High Pressure (HP) Cell Module	12

2.1.4. Mass Spectrometer Chamber	18

2.1.4. Differential Pumping System for Electron Energy Analyser	20

2.1.5. Additional Equipment	21

2.2 Part 2: Ambient Pressure Electron Energy Analyser	24

2.2.1. Ambient Pressure Electron Energy Analyser	24

2.2.2. Control system	27

2.2.3. Additional equipment	28

3.	Additional terms	30

3.1 Schedule	30

3.1.1 Final design review	31

3.1.2 Design, source code and drawings	33

3.2 Compliance with SOLARIS	35

3.3 Cooperation with SOLARIS	36

3.4 Quality control	36

3.5 Assembly Identification	36

3.6 Standard Tools	36

3.7 Packaging and delivery	36

3.8 Local staff training and remote support	37

3.9 Connecting to the backbone infrastructure	37

3.10 Elements not covered by the Contract	37

3.11 Factory Acceptance Tests (FAT)	37

3.12 Site Acceptance Tests (SAT)	39

4.	List of appendices	40

# General information

The subject of the Contract is the design, manufacturing, delivery, assembly and testing of a complete system for Near Ambient Pressure X-ray Photoelectron Spectroscopy (NAP-XPS) with high-pressure cell (HP cell). The system, hereinafter referred to as the NAP-XPS end station, will be installed on the second branch of PHELIX beamline at the National Synchrotron Radiation Centre SOLARIS in Kraków. The NAP-XPS end station will provide qualitative and quantitative information about the chemical state of the material surface as well as the chemical composition of the interface between solids, liquids and gases. The NAP-XPS end station will be dedicated to in situ and operando X-ray photoelectron spectroscopy experiments in the presence of a controlled gaseous atmosphere at pressures up to 50 mbar as well as under ultra-high-vacuum conditions, with the sample heating up to at least 800°C.

The order is divided into two parts:

Part 1 concerns the Near Ambient Pressure XPS system (hereinafter referred to as the NAP-XPS system) including a fully equipped UHV analysis chamber, a high-pressure cell as well as pumping for NAP-XPS system and differential pumping system of the electron energy analyser with required electronics.

Part 2 concerns the Ambient Pressure Electron Energy Analyser (hereinafter referred to as Analyser system) with an electrostatic lens system and digital camera-MCP detector system, capable of recording XPS at ambient pressures.

The responsibilities of the Contractor of Part 1 and the Contractor of Part 2 regarding the installation of equipment have been specified in the draft of the contract. The scope of the Contract also includes the training of SOLARIS staff in terms of safe use and basic servicing of the supplied system. Each Contractor is responsible for training within their part of the Contract.

The Contractors are required to comply with essential systems parameters specified below, boundary conditions such as the available space in the experimental hall, other components already distributed or planned in the experimental hall and the characteristics of the radiation source. The layout of the components being in the design phase on the day of the call for tender, such as the optical elements of the second branch of the PHELIX beamline, will be delivered to the Contractor of Part 1 by SOLARIS as soon as it is received, at the latest for the Startup meeting (2 weeks after signing the Contracts).

The complete floor plan with the position of the second end station, sample position in the X-ray beam, and visualization of available space for NAP-XPS end station, is attached as a separate pdf and dwg file in Appendix MECH3. The Contractor of Part 1 of the tender is required to present the concept of connecting both end stations (connection between the UHV RDS chamber from Phelix and Analysis Chamber from NAP-XPS end station) and fit into the outline presented in the floor plan (Appendix MECH3). The main Analysis Chamber of NAP-XPS system shall be adapted to the planned connection. In addition, a solution for taking the flag style holder from PTS type holder shall be presented.

SOLARIS National Synchrotron Research Centre is a third-generation light source located in Kraków, Poland. It is a low-energy facility with a storage ring based on 12 Double-Bend Achromat (DBA) cells and a circumference of 96 m. Electron bunches are injected from the linac (linear accelerator) at the energy ~550 keV, then they are ramped to the final energy of 1.5 GeV in the storage ring. Maximum number of bunches is 32, and the maximum current of the stored beam is 500 mA. The designed horizontal beam emittance is 6 nmrad, vertical is 60 pmrad. The critical energy of the synchrotron radiation beam from the bending magnet (magnetic field of 1.31 T) is approximately 1.98 keV.

In the synchrotron compact design, twelve straight sections are present, each of 3335 mm in length. Ten of them are available for different types of devices generating electromagnetic radiation (insertion devices). The PHELIX beamline utilizes the variable polarization radiation from the Elliptically Polarizing Undulator (EPU).

## 1.1. General specification of the beamline

The PHELIX beamline has the elliptically polarizing undulator (EPU) APPLE II type and a plane grating monochromator (PGM), covering the soft X-ray energy range from 50 to 2000 eV. The APPLE II undulator consists of permanent magnets with two opposite jaws that may be moved along the electron beam to create linear, circular or elliptical polarization of light. The energy resolution is limited by the resolving power of the diffraction gratings in the PGM monochromator which equals to 10 000.

The optical system of the second branch of the PHELIX beamline, despite common units for both branches like a collimating mirror (M1), a plane mirror (M2) with three gratings (G1, G2, G3) installed in the monochromator, consists of a chamber with focusing switching mirror (M3) and Exit slits unit. In addition, the optical system will consist of a refocusing setup with two mirrors installed before the NAP-XPS end station (M4a and M4b) – this part is in the design stage. Between last optical element and end station the special ion pump for differential pumping will be installed. To enable connection between the beamline and end station part, a dedicated port in the analysis chamber is required.

The PHELIX beamline, which will supply synchrotron radiation to the NAP-XPS end station is characterized by the following parameters:

a) 	Radiation source: Elliptically Polarizing Undulator (EPU), APPLE II type with permanent magnets. The undulator period length is 58 mm with a minimum gap of 14 mm;

b) 	Photon energy range: 50 – 2000 eV;

c) 	Maximum resolving power (RP) ≥ 10 000;

d) 	The maximum size of the excited area on the sample is 100 µm × 100 µm;

e) 	Photon flux on the sample ≥ 1×1013 photons/s/0.1% B.W. for 50 eV and ≥ 1×1011 photons/s/0.1% B.W. for 1500 eV, for maximum RP and electron current in the storage ring (500 mA);

f) 	Controlled polarization: variable linear, circular and elliptical.

# Scope of delivery

## 2.1 Part 1: Near Ambient Pressure XPS system

The scope of Part 1 of the order includes the following groups of elements:

a)	Analysis chamber for X-ray photoelectron spectroscopy with equipment, including:

- cryogenic 4-axes UHV manipulator,

- additional source of X-ray radiation,

- electron flood source,

- pumping system with pumps, bakeout system, power supplies, controllers and cabling;

b)	High pressure (HP) cell module with equipment, including:

- high pressure cell with manipulator,

- HP Load Lock chamber with pumping and bakeout system,

- gas manifold system;

c)	Mass spectrometer chamber;

d)	Differential pumping system for electron energy analyser with controllers and cabling;

e)	Additional equipment.

Additional requirements for the entire NAP-XPS system:

- The supplier shall ensure that the dimensions of the whole NAP-XPS system are compatible with the available space for the end station (see the floor plan attached as a separate pdf and dwg file in Appendix MECH3).

- The general geometry of the NAP-XPS system shall foresee a connection line (transfer line) to the existing UHV Radial Distribution Chamber from the PHELIX end station (first branch of PHELIX beamline). If this is deemed favourable, the supplier may foresee the use of an axial or circular distribution chamber. The concept of the transfer line has to be delivered with the offer. In addition, a solution for taking the flag style holder from PTS type holder (from UHV Radial Distribution Chamber of the PHELIX end station) shall be presented. The concept has to be implemented in a separate set of files, taking into account the floor plan and available space provided in Appendix MECH3.

- All chambers shall be equipped with a UHV pumping system. The UHV pumping system should be able to maintain UHV conditions during sample transfer between all chambers of the NAP-XPS end station. In other words, during normal operation, i.e. with all pumps working and with vacuum in all chambers at the guaranteed base level, the opening of any of the gate valves between chambers should not cause a significant (less than 5×10-9 mbar) vacuum drop in the system.

- The horizontal spot size of the sample is minimal for an incident angle on the sample of 54.7° from the surface normal, which implies that the angle between the central axis of the lens system of the ambient pressure electron energy analyser and the central axis of the beam entrance port shall be 54.7°±0.3°. The supplier may suggest a different angle with a detailed motivation.

- All turbo pumps shall be equipped with gate valves with pneumatic actuators. All pumping systems must be equipped with automatic safety valves to shut off the feed pumps in the event of power loss. Additionally, all turbo pumps should be equipped with dedicated fan for cooling (sufficient for normal operation) and with water circuit used during bakeout process.

- All vacuum piping shall be vacuum compatible and corrosion resistant.

- All chambers have to contain standard series CF mounting flanges.

- The system shall be compatible with “flag-style” sample holders with onboard thermocouple contacts for precise temperature measurements on the sample.

- Racks in a suitable quantity for all electronic units.

- The NAP-XPS system shall include an Interface for system control connected with programmable logic controller (PLC) for all necessary system components, specifically pumping system control, gas manifold system control, valves control, HP cell docking to analyser, pressure and pumping control to HP cell.

A detailed description of all groups of components with technical and functional requirements is provided below.

### 2.1.1. Analysis Chamber for X-ray Photoelectron Spectroscopy

Analysis Chamber for X-ray photoelectron spectroscopy optimized for the NAP-XPS system at SOLARIS shall be manufactured of mu-metal to enable operation over the entire energy range (50-2000 eV) and limit the influence of external magnetic fields on measurements. The residual magnetic field in the measuring area in the Analysis Chamber shall not exceed 0.5 μT. Required base pressure after system bakeout at 120 °C and after 48 hours from the bakeout shall not be higher than 5×10-10 mbar. Leak proofness test after assembly of all parts should indicate at least 1×10-12 mbar l/s. Standard series CF flanges in different sizes are required in the chamber for connecting current and future components. The detailed specification is listed below:

- The analysis chamber shall be designed to house the ambient pressure electron energy analyser in a horizontal position – one mounting flange DN200 CF.

- The analysis chamber shall be designed to house one DN63 CF mounting flange for the synchrotron radiation source (technical details shall be determined during the preliminary design stage).

- The analysis chamber shall be designed to allow the use of high-pressure cells as described in section 2.1.3.

- The analysis chamber shall be designed to use the FTIR spectrometer as described in section 2.1.3. The analysis chamber shall foresee ports for FTIR windows. The FTIR spectrometer and FTIR windows are not part of this tender. Technical details shall be determined during the preliminary design stage.

- The analysis chamber shall be designed to have a mounting flange and a gate valve for separating the Analysis Chamber of the NAP-XPS system from the connection line to the existing UHV Radial Distribution Chamber from the PHELIX end station (in accordance with the design of the Connection line provided with the offer).

- The analysis chamber shall foresee ports for a gas line connection.

- The analysis chamber shall have one mounting flange for the cryogenic 4-axes UHV manipulator, enabling translation along X, Y and Z directions and one type of rotation as described in section 2.1.2.

- The analysis chamber shall have one mounting flange for a twin-anode (Al/Mg) X-ray source.

- The analysis chamber shall have one mounting flange for X-ray monochromator.

- The analysis chamber shall have one mounting flange for a flood source.

- The analysis chamber shall have one mounting flange for a depth-profiling ion source.

- The analysis chamber shall have one mounting flange for a movable total fluorescence detector.

- The analysis chamber shall have one mounting flange for the UV source.

- The analysis chamber shall have mounting flanges for the pumping system and pressure measurement system.

- The analysis chamber shall have one mounting flange for the CCD camera.

- The analysis chamber shall have four mounting flanges for viewports providing X-ray protection, including DN40 CF and DN63 CF. One port should be located opposite the synchrotron radiation entrance.

- The analysis chamber shall have four spare mounting flanges (at least three with preferable axes passing through the analysis point), including at least one DN100 CF, and DN63 CF and DN40 CF.

Additional design requirements for the Analysis Chamber necessary to achieve measurement parameters of the Analyser system guaranteed by the Contractor of part 2 will be provided by the Contractor of part 2 within 1 month of signing the Contract. The Contractor of part 2 will send the appropriate files in electronic form to the address specified by the Ordering Party, including in the correspondence the representative of the Ordering Party indicated in the contract.

### 2.1.2. Equipment of the Analysis Chamber

The detailed specification of the analysis chamber’s equipment is listed below:

- Pumping system with pumps, power supplies, controllers and bakeable cabling. The pumping system shall ensure the base pressure not worse than 5×10-10 mbar.

The system shall contain at least the following elements:

- Turbomolecular pump, with a pumping speed ≥ 685 l/s for N2 with adequate dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h),

- Pneumatic gate valve for turbomolecular pump,

- Ion pump with a pneumatic gate valve, with a pumping speed ≥ 400 l/s for N2, with bakeout system,

- Titanium sublimation pump (TSP)

- Automatic safety and venting valves,

- Pressure measurement system from UHV to 1 atm, consisting of vacuum gauge, Bayard-Alpert gauge, Pirani-type gauge, Pirani/cold cathode, along with control electronics,

- Control unit - programmable logic controller (PLC) for pump operation with visualization of all components of the pumping system, allowing control of valves, pumps, pressure (two channels) and bakeout procedure, resistant to breakdowns, such as voltage drop (UPS protection), cooling water cut-off, etc.,

- Other essential vacuum equipment.

- Cryogenic motorized 4-axes manipulator allowing translation along X, Y and Z directions and rotation around Z-axis. All motion shall be motorized and equipped with encoders mounted at the motion axis. The manipulator should be adjusted for measuring geometry with the ambient pressure electron energy analyser in a horizontal position. The manipulator should be adjusted for sample transfer from the existing UHV Radial Distribution Chamber from the PHELIX end station (in accordance with the design of the Connection line provided with the offer).

The manipulator shall include the following elements:

- XY-motion module, motorized with a stroke length of ± 12.5 mm, total stroke square 25 mm x 25 mm, with a precise position indicator. Motors shall be equipped with encoders and limit switches. Manual mode resolution ≥ 5 μm, motorized mode resolution ≥ 1 μm.

- Z-motion module adjustable to the size of the analysis chamber, motorized with a precise position indicator. Motor shall be equipped with encoder and limit switch. Manual mode resolution ≥ 500 μm, motorized mode resolution ≥ 10 μm.

- Rotation module for rotation around the Z-axis (polar rotation), within a range of ±180°, motorized with a precise position indicator. Motor shall be equipped with encoder and limit switch. Motorized mode resolution of 0.1°.

- The manipulator shall have two sample stations (holder seats) compatible with flag-style sample holders. Both sample stations have to be placed around the Z-axis with movement in the same axes.

- The first sample station shall be optimized for cooling mode with liquid helium. Liquid helium cooling unit shall be equipped with necessary silicon diodes, and an inlet adapter for connecting the helium levar. A helium transfer line shall be included. It shall allow to cool down the sample holder to a temperature of 10 K or less with liquid helium in the open circuit and heat up to at least 470 K. Temperature should be controlled with a temperature accuracy of ±1 K. The supplier shall provide a cooling profile for liquid helium cooling, which specifies temperature as a function of time and the obtainable minimum temperature.

- The second sample station shall be optimized for heating mode. It shall be equipped with a set of mechanical and electrical feedthroughs and connections as well as thermocouple type K. In addition, temperature measurement shall be on the transferable sample holder using a type K thermocouple, i.e. contacts need to be provided. Sample heating shall be achieved separately by resistive heating and e-beam heating. It shall allow sample heating to at least 500 °C in the case of a resistive heating system and 1200 °C in the case of an e-beam heating system in UHV conditions. The temperature stability shall be better than 1 °C. The temperature regulation shall be PID-controlled to avoid overshooting by more than 5 °C. Heating the sample to a higher temperature should not cause any disturbance in measurements.

- It shall be possible to measure the drain current on the sample for the first sample station (station for cooling mode). Further, the application of a low to medium voltage bias (0 to ~ 200 V) shall be possible. Dedicated electrical feedthrough and connections should be foreseen for drain current measurements.

- Power supplies for heating and cooling sample stations equipped with PID temperature controllers (separate for each sample station), cables and necessary connections for resistive and electron heating. Manual and PID control of the temperature shall be possible.

- Reference sample – Au polycrystalline plate mounted permanently.

- Additional source of X-ray radiation:

- The X-ray source shall be compatible with the vacuum system, with the bake-out temperature ≥ 200ºC.

- Twin anode (Al/Mg) X-ray source shall be installed with a power supply and all necessary cabling. The maximum power for twin anode shall be 600 W for Al anode and 400 W for Mg anode. The anode voltage shall be in the range of 0.5 – 15 kV, with 0.1 kV accuracy.

- The X-ray source shall be closed-circuit water-cooled.

- The X-ray anode shall be possible to align, with a linear shift of at least 70 mm and a tilt range of ±4°.

- Microprocessor-based emission controller with graphical interface and possibility of communication and control from PC.

- System for equipment and people protection: flow, temperature and cooling water pressure monitoring, vacuum drop protection, protection against removing the shield of the lamp while operating, etc.

- Electron flood source for charge neutralization on the sample surface:

- The flood source shall be installed with a power supply and all necessary cabling.

- The flood source shall be compatible with the vacuum system, with the bake-out temperature ≥ 200ºC.

- The flood source shall operate in the minimum energy range of 0.01–1000 eV. The flood source shall operate in two modes: low-energy (0.01-50 eV; with a resolution of 0.01 eV) and high-energy (50-1000 eV; with a resolution of 0.1 eV).

- The flood source shall operate with an electron emission current from 1 to 1000 μA with a resolution of 0.2 μA.

- Laser source for choosing an analysis spot on a sample.

- Four viewports with protection from X-ray radiation shall be included.

- At least two cameras for sample transfer and positioning within the HP cell: one ultra-zoom camera (zoom x100) with a monitor. Holders, cabling and software with a license for all cameras shall be included.

- Blanks for spare flanges.

### 2.1.3. High Pressure (HP) Cell Module

High pressure (HP) cell module shall be optimized for in situ/operando X-ray photoelectron spectroscopy experiments over the entire energy range (50-2000 eV) and in the presence of a controlled gaseous atmosphere in the pressure range from 1 ×10-7 mbar to 50 mbar, with the sample heating up to at least 800 °C. The Analyser System and the Analysis Chamber for X-ray photoelectron spectroscopy shall be compatible with the use of exchangeable high pressure cells. In the present tender, one high pressure cell defined below shall be included. The detailed specification of HP Cell module with components is provided below:

- High pressure (HP) cell will be used in reaction studies with well-defined conditions, with special emphasis on high pressure experiments. The cell shall be designed to perform in situ XPS measurements up to at least 50 mbar. The detailed specification is listed below:

- The HP cell, including the pumping line, shall be made of corrosion-resistant materials. The supplier shall specify the materials’ properties and to which extent these materials are corrosion-resistant.

- The interface between the high pressure cell and the analyser shall be constructed so that the pressure in the analysis chamber does not exceed 2 × 10-6 mbar when the high pressure cell is operated at 30 mbar.

- The HP cell shall allow sample transport to the Analysis Chamber without losing UHV conditions.

- The HP cell shall allow measurements in two modes: static with pumping through the electron energy analyser and dynamic with additional pumping through the HP cell.

- The HP cell shall contain two sets of front covers with specific HP windows for: synchrotron radiation, X-ray beam (dual anode), monochromated X-ray beam (dual anode), IR beam inlet and outlet, UV light, ultra-zoom camera and LED light. The detailed specification of front covers with specific HP windows shall be determined during the final design stage. The design of the HP cell shall allow easy replacement of the windows, without Analysis Chamber venting. The supplier shall specify the procedure to replace the windows.

- The synchrotron light shall enter the cell through an X-ray entrance window. The window shall be replaceable. The window shall withstand a pressure differential of at least 60 mbar. The supplier shall specify the X-ray transmission of the window supplied with the cell in the energy range of 50 eV to 2000 eV.

- The synchrotron light shall not be screened from reaching the accurate focus point of the analyser on the sample.

- The entrance aperture to the electron energy analyser – the so-called analyser first cone shall be mounted on the HP cell. The design of the HP cell shall allow easy replacement of the aperture. The supplier shall specify the procedure to replace the apertures.

- Accurate positioning of the sample within the HP cell is necessary. Experiments in high pressures require very small distances between the sample and aperture (entrance aperture to the electron energy analyser). Typically the distance between the sample and aperture (d label shown in scheme 1) will be twice the opening diameter of the aperture. Since apertures with sizes down to 100 μm will be used, the position relative to the aperture shall be controlled to within 2 μm.

- The HP cell shall be compatible with flag-style sample holders. It shall allow the use of solid samples with a maximum size of 1x1 cm2 and a thickness at least in the range of 0.5 – 4.0 mm.

- The HP cell shall foresee heating by laser illumination from the back side of the sample holder. A sample temperature of at least 800°C shall be reachable within the cell in the UHV conditions and 30 mbar of N2. The temperature stability shall be better than 1°C. The temperature regulation shall be PID controlled to avoid overshooting by more than 5°C. In addition, manual control shall be possible.

- The HP cell shall be equipped with a pressure measurement system, and all necessary gauges (vacuum gauge, Pirani-type gauge, Pirani/cold cathode) with controllers. Pressures in the range from the maximum pressure down to 1 x 10-7 mbar shall be measurable. Care shall be taken to measure the pressure as close to the surface of the sample as possible. The HP cell pressure shall be measured directly on the cell using a Pirani gauge.

- The HP cell shall include infrared transparent windows (FTIR windows) allowing for optical access to the sample in a grazing (approximately 10 degree) reflection geometry (see scheme 1).

- The NAP-XPS system and HP cell shall foresee the geometry and space needs for an FTIR spectrometer, detector and relevant beam optics. The beam path shall be kept to a minimum. The FTIR spectrometer, detector and beam optics are not part of this tender. The detailed specification shall be determined during the final design stage.

- The HP cell shall be made from appropriate materials to avoid contamination of the interior surfaces. It is not allowed to use materials with molybdenum and tantalum in the HP cell. Materials used to produce the HP cell shall be approved by the SOLARIS during the final design stage.

- The gas inlet position and geometry of the cell shall be designed to maximise the relative local pressure at the point of measurement.

- The gas inlet and outlet lines shall be configured to allow analysis by complementary gas-phase techniques, such as mass spectrometry.

- The HP cell shall be removable from the Analysis Chamber to allow measurements in UHV conditions.

- The HP cell housing and receiving station shall be water-cooled as appropriate for the chosen heating concept and geometry. Additionally, the front cone of the Analyser system shall be water-cooled to prevent the Analyser from heating up when the sample is heated at higher temperatures > 800 °C.

- The HP cell shall be electrically isolated and it shall be possible to measure the potential on the sample.

- It shall be possible to transfer sample holders with a transfer arm (from the second end station through the designed connection line) directly to the HP cell without using the wobble stick.

- The HP cell shall have a reliable, easy-to-use and tight mechanism (“the door”) separating the HP cell volume from the Analysis Chamber. The door operation shall be automated.

- Docking of the HP cell to the Analyser shall be fully automated and shall use a single motor with docking force control.

- The HP cell shall have two ways/routes for pumping the HP cell volume, with high and low pumping rates. Both routes shall not add any volume to the HP cell. Both routes shall be equipped with in-vacuum, automated valves directly at the exit of the HP cell volume.

- Manual and PID control of the temperature shall be possible. The temperature controller shall have a “device server” enabling connection to the SOLARIS control system, thanks to which it will be possible to read and control the sample temperature.

Scheme 1. Schematic drawing of aperture-sample geometry with X-ray and IR beams within high pressure cell.

- High pressure cell manipulator shall allow for sample heating by laser illumination and water cooling. The detailed specification is listed below:

- The HP cell manipulator shall be mounted on a special trolley, which shall allow easy movement of the HP cell with the manipulator into/from the Analysis Chamber.

- The HP cell manipulator shall be compatible with “flag-style” sample holders.

- The HP cell manipulator shall allow control of at least three axes: x, y, and z. All motions shall be motorised.

- XY-motion module, motorized with a stroke length of ± 5.0 mm with a precise position indicator. Motors shall be equipped with encoders and limit switches. Manual mode resolution ≥ 5 μm, motorized mode resolution ≥ 1 μm.

- Double Z-motion module, motorised with movement in double Z-axes. The stroke length of the z motion shall be determined based on the vacuum system geometry. The main Z-motion module shall be used for HP cell proper placement in the analysis chamber. The second Z-motion module shall be used for highly precise sample positioning in the HP cell.

- The manipulator shall allow for sample heating with an integrated thermocouple and temperature controller. The supplier shall provide a heating profile for laser heating, which specifies temperature as a function of time and the obtainable maximum temperature.

- In the design of the sample holder station and heaters, compactness shall be aimed to reduce degassing rates.

- It shall be possible to remove the HP cell manipulator with HP cell from the NAP-XPS system without disassembling any vacuum parts of the cell.

- HP Load Lock chamber shall be included in the HP cell module to facilitate sample loading directly into the HP cell from the air side. The detailed specification is listed below:

- The HP Load Lock shall be designed so that fast sample holder loading into the HP cell is possible.

- The HP Load Lock shall include a gate valve toward the Analysis Chamber.

- The HP Load Lock shall reach a pressure of less than 9 × 10-7 mbar within 2 hours after pump down.

- The HP Load Lock shall be equipped with the necessary pumping system.

- The pumping system shall include a turbomolecular pump, with a pumping speed ≥ 300 l/s for N2 with adequate dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h). A pneumatic gate valve for the turbomolecular pump shall be foreseen.

- The pumping system shall include a dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h), dedicated for HP cell pumping with necessary valves.

- The pumping system shall include a pressure measurement system from UHV to 1 atm, along with control electronics.

- The pumping system shall be automated and prevent damage to the HP cell windows during the pumping.

- The pumping system shall include microprocessor control unit for pumps operation with visualization of all components of the pumping system, allowing control of valves, pumps, pressure and bakeout procedure, resistant to breakdowns, such as voltage drop (UPS protection, interlocks etc).

- The pumping system shall include automatic safety and venting valves, tubing and wiring, power supplies, controllers and bakeable cabling.

- Gas manifold system shall be designed to deliver controlled flow, pressure and composition of gases to the HP cell. The detailed specification is listed below:

- The gas manifold system shall consist of two separate lines delivering gas directly to the HP cell. In-vacuum valve on the HP cell separating the two gas lines shall be included.

- The gas manifold system for the HP cell shall cover the entire range of pressures from 10-7 mbar to 50 mbar. It shall be possible to regulate the pressure using the flow controllers and the pressure regulator.

- It shall be possible to close the flow controllers by placing the valve directly after each flow controller.

- The gas manifold system shall include gas delivery and exhaust lines to and from the HP cell.

- The gas manifold system shall be equipped with safety valves that will open if the pressure gets too high, for protection of the windows etc. in the connected chamber.

- The gas delivery and exhaust lines to the HP cell shall foresee connection to the mass spectrometer by leak valves. The leak valve shall be located directly on the gas line. It shall be possible to switch automatically between probing the inlet and outlet gas from the HP cell with fixed leak rates to the mass spectrometer.

- The gas manifold system shall be connected to 4 gas lines for four different gases, with a simple gas mixing scheme (provided during the design stage by SOLARIS) and four mass flow controllers (MFC). The decision on which MFC shall be connected to a given manifold and on the type of gas used in the system will be communicated to the Contractor by SOLARIS during the Startup meeting.

- All fittings shall be VCR types or equivalent (with replaceable gaskets).

- The design of the gas manifold system shall foresee the future extension of the system with additional gas lines as well as the installation of filters and cold traps located before the flow controllers.

- Control unit - programmable logic controller (PLC) shall be included for control of the gas system, operation and visualization of all components. Interface for system control shall be included.

### 2.1.4. Mass Spectrometer Chamber

The mass spectrometer to be mounted on the mass spectrometer chamber will be tendered in a separate procurement process. The procurement will be finalised prior to the design phase of the present NAP-XPS system, so that all mass spectrometer dimensions and specifications will be available in the design phase of the mass spectrometer chamber. The dimensions of the mass spectrometer chamber shall be defined in collaboration with the SOLARIS in the design phase. The detailed specification of the mass spectrometer chamber is listed below:

- The mass spectrometer chamber shall be located as close as possible to the HP cell manipulator and the gas delivery and exhaust lines.

- The connection flange for the mass spectrometer shall be DN 63 CF.

- The chamber shall have a leak valve.

- The mass spectrometer chamber and pumping system shall be designed for mass spectrometer operation, taking into account maximum pressure operation of the spectrometer: 5×10-6 mbar (option for 1×10-4 mbar).

- The HP cell’s inlet and outlet lines as well as the first differential pumping stage of the analyser shall be connected to a mass spectrometer via a combination of automatic angle and leak valves, providing a way for the quick analysis of the gas composition before or after its contact with the sample. The design shall avoid artificial pressure build-ups and dead volumes in the gas pipes. The system shall allow change from probing the inlet to probing the outlet gas composition within 5 sec.

- The mass spectrometer chamber shall be connected to a differential pumping system. The pumping system shall include a turbomolecular pump, with a pumping speed ≥ 67 l/s for N2 with an adequate dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h). A pneumatic gate valve for the turbomolecular pump shall be foreseen.

- The pumping system shall include a pressure measurement system from 
1x10-8 mbar to 1 atm, along with control electronics.

- The pumping system shall include microprocessor control unit for pumps operation with visualization of all components of the pumping system, allowing control of valves, pumps, pressure and bakeout procedure, resistant to breakdowns, such as voltage drop (UPS protection, interlocks etc).

- The pumping system shall include automatic safety and venting valves, tubing and wiring, power supplies, controllers and bakeable cabling.

### 2.1.4. Differential Pumping System for Electron Energy Analyser

The differential pumping system for the ambient pressure electron energy analyser shall consist of at least three pumping stages with all necessary pumps, pressure measurement system and accessories, which enables analyser to work in the requested pressure range from UHV up to 50 mbar. The detailed specification of the analyser pumping system with components is provided below:

- 1st stage of pumping system for analyser pre-lens. It shall include two turbomolecular pumps, with a pumping speed ≥ 300 l/s for N2, with adequate dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h). Pneumatic gate valves for the turbomolecular pumps shall be foreseen. It shall also include automatic safety and venting valves, tubing and wiring, power supplies and controllers.

- 2nd stage of pumping system for analyser lens. It shall include one turbomolecular pump, with a pumping speed ≥ 300 l/s for N2, with adequate dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h). An electro-pneumatic gate valve for the turbomolecular pump shall be foreseen. It shall also include automatic safety and venting valves, tubing and wiring, power supplies and controllers.

- 3rd stage of pumping system for analyser lens/hemisphere. It shall include one turbomolecular pump, with a pumping speed ≥ 300 l/s for N2, with adequate dry oil-free and fluorine-free fore-vacuum pump (with a pumping speed ≥ 15 m3/h). An electro-pneumatic gate valve for the turbomolecular pump shall be foreseen. It shall also include automatic safety and venting valves, tubing and wiring, power supplies and controllers.

- All pumping stages shall be equipped with pressure measurement systems from UHV to 1 atm, consisting of UHV gauges, Pirani-type gauges, pirani/cold cathode gauges along with control electronics.

- The pumping system shall include the bypass between the Analytical Chamber and analyser with tubing and safety valves.

- Microprocessor control unit for operation and visualization of all components of the pumping system, allowing control of fore-vacuum pumps, turbomolecular pumps, valves, and pressure measurement system. It shall include necessary interlocks to protect the vacuum system during a power failure or in the case of a breakdown of its components.

- Other essential vacuum equipment.

### 2.1.5. Additional Equipment

The following items have to be compatible with the UHV system and fulfil the required performance parameters:

- Modular instrument frame: The whole instrument shall be mounted on the adjustable, modular, rigid, steel frame. The modular frame shall allow for easy disassembly of the HP cell module (HP cell module mounted on the special trolley). The modular frame shall allow for easy disassembly of the Analyser system. Analyser system, along with the differential pumping system shall be mounted on a special trolley, which shall enable its disassembly and removal from the Analysis Chamber. It shall be possible to adjust the position of the entire instrument (analysis chamber, HP cell and electron analyser) relative to the radiation beam (at least ±20 mm in the X, Y axis and at least ±25 mm in the Z axis range). It has to be adapted to system bake-out and equipped with a stainless steel mounting plate and insulating plate. Air and water distribution systems shall be included.

- Bake-out system with temperature control and the possibility of independent bake-out of all chambers in the system (separate for analysis chamber, HP cell module). Individual heating zones have to be separated from each other and from the surroundings by thermal curtains. The bake-out has to be done by warm air supplied by the air blow system, controlled by a microprocessor temperature controller, with a temperature increase of at least in the range of 1°/h to 50°/h. The bake-out system shall be equipped with sensor break protection and additional protection in the bake-out zone in case of thermocouple failure. It must be possible to define duration of the bakeout in the range from 1 min to at least 250 hours with 1 min accuracy.

- Flag style sample holders for transferring samples within the UHV system (at least 10), of which 4 holders with onboard thermocouple contacts for precise temperature measurements on the sample. For use in the HP cell, holders shall be corrosion-resistant and made of chemical inert materials. The material will be defined during design period.

- Racks (19" cabinet) in a suitable quantity for all electronic units. The electrical connection shall be protected by a main switch and an emergency stop. All connection cables shall be included. It shall be possible to connect particular electronic units to the UPS power supplies available at SOLARIS. At least two power supply circuits shall be envisaged and prefabricated for delivered racks with electronics. One for pumping systems, including pumps, vacuum gauges, PLC units, controllers along with controllers for X-ray lamp, flood gun etc. The second circuit shall be intended for the bakeout system with bakeout power supplies and controllers. The detailed electrical specifications of the SOLARIS standards and units are described in “Appendix EL1”. All details will be discussed during design phase.

- LED illumination system inside the chambers.

- A chiller system with a heat exchanger for closed-circuit water cooling shall be included along with necessary flow meters and cables. One chiller system for all devices is preferred.

- Infrastructure of the UHV system (power and low voltage electrical installations, water and compressed air systems, etc.) between the system components and the backbone infrastructure of SOLARIS.

- Interface for system control and automation. The software with component modules shall allow operation and control of the entire system, i.e. analysis chamber, manipulator, HP cell, sample temperature, pumping system, gas manifold system, valves, bake-out equipment.

- Set of manuals, installation, connecting the UHV system infrastructure and training. The technical documentation for all components provided by the Contractor has to be sent to SOLARIS on the day of delivery. It includes a description of technical parameters, pumping and heating procedures, manuals for regular usage, a description of assembly/disassembly (if regular usage so requires), a description of regular usage and maintenance. The contractor has to provide training for a maximum of 8 people, with a total duration of up to 32 clock hours. The training has to be carried out during the installation of the object of the Contract at SOLARIS.

- All motors and encoders (motion modules) for Cryogenic 4-axes manipulator and HP cell module (without motion module responsible for cell docking) shall be compatible with the IcePAP motion control system. Device server” in Tango Controls 9.2+ with stepper controllers complying with SOLARIS standards (IcePap) will be used.

- For all components and subsystems original Company software should be provided. Software should allow for control and standard operation of all components of NAP-XPS system. Additionally, “Device servers” in Tango Controls 9.2+ for the main components of the system should be provided: heating and cooling controller for manipulator and HP cell, additional X-ray source, electron flood source, control of the pumping system (i.e. for HP cell module, HP load lock, Analysis chamber, mass spectrometer chamber, energy analyser). The contractor will provide technical support in the field of software updates and implementation of additional functionalities (e.g. additional commands, attributes) of Device servers over a period of 36 months, starting from the completion of the stage mentioned in § 4, section 1.3 of Appendix 2 to the ToR, in the total amount of 24 working hours.

- SOLARIS will provide a network switch to handle the network traffic of devices and will impose network addressing in accordance with SOLARIS standards, enabling, among others, making copies (backup). Access to the Administrator account on the server is required. Details will be discussed with the Contractor during the first meeting (Starting Meeting).

## 2.2 Part 2: Ambient Pressure Electron Energy Analyser

The scope of part 2 includes the design, manufacturing and installation of an Ambient Pressure Electron Energy Analyser with combined capability of Ambient Pressure X-ray Photoelectron Spectroscopy and Imaging X-ray photoelectron spectroscopy. The Analyser system shall feature a hemispherical electron energy analyser amended with a separate pre-lens that provides efficient differential pumping as well as electron optics for imaging XPS. The order includes:

- Ambient pressure electron energy analyser,

- Control system,

- Additional equipment.

A differential pumping system for the Analyser system with necessary cabling, connections, safety valves, venting valves and microprocessor control unit for pumps will be provided by the Contractor of part 1 of the tender. A detailed description of all groups of components with technical and functional requirements is provided below.

### 2.2.1. Ambient Pressure Electron Energy Analyser

The hemispherical electron energy analyser with combined capability of Ambient Pressure X-ray photoelectron spectroscopy and Imaging X-ray photoelectron spectroscopy. The Analyser system should be optimised for quick high-resolution photoemission spectroscopy (PES) measurements for high throughput data collection. The shielding of electron energy analyser shall be manufactured of mu-metal (mu-metal plus mu-metal liner) to enable operation over the entire energy range (50-2000 eV) and limit the influence of external magnetic fields on measurements. The Analyser system shall be equipped with at least two front cone sets with a different opening diameter (entrance aperture to the electron energy analyser) to enable experiments from UHV to up to 50 mbar (depending on the used cone). The detailed specification is listed below:

- Hemispherical electrostatic energy analyser of charged particles with a mean radius of a minimum 200 mm and with high transmission. The analyser shall be equipped with a double mu-metal shielding (mu-metal plus mu-metal liner) to allow operation over the entire energy range (50-2000 eV) and limit the influence of external magnetic fields on measurements. It shall be possible to connect a minimum of 5 turbomolecular pumps, mountable on DN 100CF flanges.

- The electron energy analyser shall be amended with a separate pre-lens that provides efficient differential pumping as well as electron optics for imaging X-ray photoelectron spectroscopy.

- The analyser shall have a kinetic energy range of at least 50 eV to 2000 eV in transmission mode. The analyser shall have a kinetic energy range of at least 50 eV to 1500 eV in imaging mode The analyser shall have a kinetic energy range of 100 eV to at least 1500 eV in angular mode. The supplier shall specify the kinetic energy range of the analyser.

- The energy resolution of the analyser shall be better than 15 meV FWHM at 20 eV pass energy and 500 eV kinetic energy.

- The analyser shall be equipped with a minimum of nine interchangeable slits for optimal balance between best intensity and best energy resolution. The slits shall be fully motorized and controlled by provided software. The standard setup of slits shall include: in millimetres 0.1s, 0.2s, 0.3s, 0.5s, 0.8s, 1.0s, 1.5s, 2.5s, 4.0s (s = straight).

- At least five values of pass energy shall be available in the range of 5–500 eV. The standard setup of the pass energy shall include 20, 50, 100, 200, and 500 eV.

- The supplier shall specify the electron energy analyser focus size and shape for different analyser modes. In particular, it should be specified how much the area outside of the focus contributes to the total signal, assuming homogeneous sample irradiation with a light spot much larger than the analyser focus. This contribution shall not be more than 25% of the total signal.

- The analyser shall be equipped with a high-performance digital 2-dimensional camera-microchannel plate detector (CCD/MCP) system for computer data acquisition. A low noise level is required with < 0.01 cps/channel. The detailed specification for the performance of the digital camera-MCP detector system is listed below:

- 1000 energy channels available simultaneously,

- 750 angular channels available simultaneously,

- Total noise counts for all energy channels: < 2 cps,

- Average noise per channel: < 0.01 cps,

- Available dynamic range per pixel: > 7 bit,

- Detector information read-out rate: > 70 frames per second,

- Detector modes: pulsed and ADC.

- It shall be possible to operate the analyser in imaging (spatial resolved) mode with a spatial resolution better than 10 μm and a minimum field of view of 0.5 mm.

- It shall be possible to operate the analyser in both transmission and angle resolved modes. The acceptance angle for angular resolved modes shall be at least ± 9°. The supplier shall describe all analyser modes in detail including a description of the modes’ use.

- In the transmission mode it shall be possible to operate analyser in the swept and fixed (so-called snapshot) mode. In the fixed mode it shall be possible to record a spectrum within seconds by taking a snap shot of the detector image, covering at least 8% of the pass energy.

- The analyser shall achieve a pressure differential between the sample environment and the detector so that pressure of 10-6 mbar or less at the detector can be maintained when the sample is at 50 mbar. Contractor of Part 2 must provide requirements to achieve that vacuum level in terms of pumping speed for each pumping stage of the energy analyser.

- The pre-lens shall be equipped with a small front aperture that permits electrons to enter the analyser while providing efficient differential pumping. The front cone angle shall allow easy photon access and flexibility for the sample handling. The front cones shall be made out of material with good chemical and thermal resistance as well as mechanical stability. The supplier shall also specify the front cone material and its properties i.e. chemical and thermal resistance. In addition to the front cone aperture, a solution shall be foreseen to enable high spatial resolution (better than 10 μm).

- The front cone shall be interchangeable to increase or decrease the aperture size. The minimum of two front cone sets has to be provided with different aperture diameters: 0.2 mm and 0.5 mm.

- The analyser shall offer a mode enhancing the transmission of electrons in the lens without applying a bias to the sample. At the same time a field-free zone between sample and lens entrance is requested, to avoid sample and position-dependent local field variations influencing results. The supplier shall state the signal increase as compared to the traditional lens mode.

- The supplier shall state the transmission losses at pressures of 1 mbar, 20 mbar and 50 mbar in comparison to the transmission in vacuum for different apertures sizes 0.2 mm, 0.5 mm and 0.8 mm, assuming a working distance of 0–2 mm and a nitrogen gas atmosphere at the sample.

- Working distance (between the sample and the entrance to the lens system) shall be at least in the range of 0.3 – 2.0 mm.

- The analyser shall be equipped with real time read-out (display count rate in real-time in the software) and electronic analyser lens deflectors for diagnostics and experimental optimisation purposes, i.e. to change the optimum position of the analyser field of view.

- The analyser shall be compatible with the use of an HP cell (described in section 2.1.3.).

- The analyser shall be equipped with mu-metal conical connection set, for proper connection between the Analyser and the Analysis Chamber.

- The analyser shall be supported on the modular frame, which shall be provided by the Contractor of Part 1 (described in section 2.1.5). Contractor of Part 2 should provide sufficient information related to supporting points of electron energy analyser together with 3D model of the outer shape to Contractor of Part 1 within 4 weeks after signing the contract.

- Contractor of Part 2 has to provide sufficient information set and documentation of used cone shape for Contractor of Part 1, to match the analyser cone with the HP cell (docking tightness). All details will be defined during the design phase.

### 2.2.2. Control system

System designed to control acquisition of photoelectron spectra with ambient pressure electron energy analyser. The Supplier shall offer the system with a graphical user interface (GUI), which will allow controlling analyser (e.g. set the lens mode, slit type, kinetic energy and pass energies) and detector (e.g. adjust the detector voltages). The supplier’s software application shall support multiple modes of measurement, at least Fixed energy (Snapshot) mode and Sweep energy mode. The software shall allow for the alignment process while searching for optimised geometrical positions.

The software shall allow for data control and acquisition (including fast acquisition mode) as well as live visualization of data. The software shall foresee the integration of additional equipment and shall allow the simultaneous recording of external signals such as temperature, current etc. The software producer shall ensure access to program updates for at least 60 months and provide program support throughout the entire period of the quality warranty. Files with collected data shall be compatible with the software used for data analysis i.e. CasaXPS, Igor Pro software. One licence for the data analysis package Igor Pro from WaveMetrics shall be included. The supplier’s software application shall allow the automatic numbering of spectra with a user-specified base name and a consecutive id number.

It shall be possible to integrate the software into the SOLARIS control system (the Tango Controls 9.2+ system). The contractor shall provide the API (Application Programming Interface) documentation for the Windows 11 operating system (64-bit) to the program for controlling, collecting and storing data, simultaneously controlling the energy analyser and detector, along with examples of API usage. The contractor shall provide an additional copy of the software to be integrated with the Tango Controls 9.2+ system and for testing purposes.

### 2.2.3. Additional equipment

The following items have to fulfil the required performance parameters:

- Mounting of standard vacuum components, and dedicated ports for differential pumping system. A differential pumping system for the analyser with necessary cabling, connections, safety valves, venting valves and microprocessor control unit for pumps will be provided by the Contractor of part 1 of the tender.

- Electronics for the ambient pressure electron energy analyser:

- Power supply system and cabling enabling the spectrometric analysis of negatively charged particles.

- The electronics of the analyser shall allow operation in fast acquisition mode.

- The high-voltage electronics shall provide ultra-stable conditions for reliable accurate measurements.

- The standard 19” racks (with air cooling) for mounting the electronics. Racks power supply guidelines, UPS, grid system, grounding and equipotential bonding requirements, low voltage switchgears requirements shall also be provided. The electrical guidelines shall be prepared taking into account “Appendix –EL1”. It shall be possible to connect particular electronic units to the UPS power supplies available at SOLARIS. At least two power supply circuits shall be envisaged and prefabricated for delivered rack with electronics. One (the UPS power) for crucial components and the second circuit (standard power) for not critical elements. All will be discussed during design phase.

- Desktop computer with the minimum parameters specified below:

- English version of Windows 11 Pro (64-bit) operating system, along with a recovery medium "system recovery" and with a carrier with drivers,

- RAM memory: 128 GB,

- Hard drives operating in RAID 1: 2 x 1 TB SSD. In the event of a hard disk malfunction, it shall stay with the Customer.

- Multi-core processor with integrated graphics, achieving in the test PassMark - CPU Benchmarks result min. 37 000 points (results available at https://www.cpubenchmark.net/cpu_list.php),

- Two 27" monitors,

- Graphics card adapted for simultaneous control of two monitors, achieving in the PassMark Performance Test result of at least 6500 points in the G3D Rating, the result available at: https://www.videocardbenchmark.net/gpu_list.php,

- 10/100/1000 Ethernet RJ 45 network card, integrated with the motherboard, supporting WoL support (user-activated function),

- 2 USB 2.0 ports and 3 USB 3.2 ports, 2 USB Type-C ports

- Quality warranty: monitors – 24 months, computer equipment – manufacturer’s quality warranty of at least 24 months in the workplace of the computer, i.e. Next Business Day,

- Equipment necessary to connect the analyser, in addition to those listed above.

# Additional terms

## 3.1 Schedule

The Contractor of Part 1 of the tender will assume responsibility for ensuring full compatibility (in particular the connection of the analyser with analysis chamber and HP cell) of both components of the end station (Part 1 and Part 2), while maintaining the performance parameters of both components declared by the Contractors. The Contractor of Part 2 will guarantee that all essential technical data and information needed to integrate both parts are provided. The Contractor of Part 2 will install the Analyser system in cooperation with the Contractor of Part 1 in accordance with the draft of the contract.

The Contractors of Part 1 and Part 2 shall present a detailed schedule of the Contract implementation and deliver it within two weeks of the date of signing the Contract.

- The schedule has to include stages: design, manufacturing, testing, delivery and installation with tests. The schedule is intended to give an overview of the design and production process and allows for easy communication on the project. The schedule may be completed during the startup meeting which will take place at SOLARIS or remotely via dedicated communicator platform two weeks after signing the Contract.

- The schedule shall contain delivery dates for documentation prior to milestones and the deadline for the NAP-XPS system delivery (Part 1) and ambient pressure electron energy analyser delivery (Part 2) as well as on-site installation and integration of all components.

| Milestones | Points to be fulfilled before the milestone is considered completed | Deadlines |
| --- | --- | --- |
| Startup meeting | The schedule. The Contractors have to recognize all technical issues and propose technical solutions. | 2 weeks after signing the Contract |
| FDR (Final Design Review) | The completion of detailed design. After FDR acceptance, SOLARIS gives permission for manufacturing. | 3 months after signing the Contract |
| FAT (Factory Acceptance Test) | The agreed tests will be performed by the Contractors and approved by SOLARIS. | To be agreed during FDR |
| Delivery of the NAP-XPS system (Part 1) | The NAP-XPS system will be delivered to the indicated address and checked for possible damage. | 11 months after signing the Contract |
| Delivery of the ambient pressure electron energy analyser (Part 2) | The ambient pressure electron energy analyser will be delivered to the indicated address and checked for possible damage. | 11 months after signing the Contract |
| On-site installation of all system components | All components will be installed in accordance with the terms of the Contract | 12 months after signing the Contract |
| SAT (Site Acceptance Tests) | The agreed tests will be performed by the Contractors. The delivered equipment has to pass all required tests. | 12 months after signing the Contract |
| Trainings | Local staff will be trained by the Contractors in the proper maintenance and safe operation of the delivered equipment. | 12 months after signing the Contract |

### 3.1.1 Final design review

A Final Design Review (FDR) has to be approved by SOLARIS, which will take place within two weeks after delivery of all necessary documents to the Ordering Party.

Final Design shall contain:

- Detailed technical description of the complete NAP-XPS end station, including integration between Part 1: Near Ambient Pressure X-ray Photoelectron Spectroscopy (NAP-XPS) system and Part 2: Ambient Pressure Electron Energy Analyser. The Contractor of Part 2 will guarantee that all essential technical data and information needed to integrate both parts are provided.

- 2D and 3D models and drawings of the complete NAP-XPS end station, including the location of all components with the defined position of all required connections (water, compressed air, power, cable connections, control panels, etc.).

- The detailed scheme of the electrical wiring and power supply.

- The detailed cooling system diagram for all components requiring water cooling.

- The detailed compressed air system diagram.

- Final list for the motorized axes (number and type of motors for motorized axes, number and type of required encoders).

- The list of the installation work.

The Contractors of Part 1 and Part 2 have to enable witness points during the assembly of critical parts. FAT in the presence of observers from SOLARIS and SAT shall be foreseen. FAT and SAT scopes will be agreed upon at the startup meeting or the FDR meeting at the latest.

The Contractors shall submit the schedule of the FAT and the SAT. The Contractor of part 1 will be responsible for coordinating the process of installing Part 1 and Part 2 so that both parts of the SAT will be held for the completely assembled NAP-XPS end station. The Contractors will agree on the tests order, taking into account the interdependence between subjects of the order, respecting the time that both parties have to devote to SATs.

If one of the two Contractors fails to meet the delivery and installation deadlines, which will result in delaying the final tests of the entire device, the formal consequences will not apply to the Contractor who met the deadlines, in accordance with the draft of the contract.

During Factory Acceptance Tests all standard tests shall be performed (vacuum tests, hydraulic function tests, mechanical metrology on all motorized motions, all electrical components shall be tested). The Contractor has to provide hardware and software needed to perform all motion tests to be carried out during the FAT. In the case of motorized axes, the Contractor shall perform motion tests using original company controllers and additionally during SAT with IcePap controllers provided by SOLARIS and together with SOLARIS engineers. The tests of all electrical components (switches, temperature sensors, etc.) and mechanical elements (modules of movement and rotation) shall be performed during the FAT. A resolution test of the analyser module shall be performed during the FAT.

The Contractor of Part 1 shall carry out the installation and the SAT of the NAP-XPS system. The Contractor of Part 2 shall carry out the installation and the SAT of the ambient pressure electron energy analyser. All necessary consumables as necessary for the end station assembly (screws, washers, caps, seals, connectors, required cooling water and compressed air circuits prefabricated components, etc.) and all other nonstandard tools will be delivered by Contractors. If the supplied equipment does not meet the specified technical requirements, the fault will be corrected immediately by the Contractor. If the fault prevents the SAT from being successfully performed, the tests will be repeated after correcting the fault. Each subsequent event of this type will result in a repetition of the procedure.

The documentation describing safety procedures during the assembly and disassembly of each subassembly shall be provided to SOLARIS not later than one month before the delivery of the subject of the order.

All communication shall be in English or Polish. The Contractors shall designate project managers, who will be the main contact persons for SOLARIS. Meetings can be held at SOLARIS, at the Contractor’s business office or online. Both parties will cover their own travel and accommodation costs, which will not be reimbursed by the other party.

The payment terms have been specified in Appendix 2 to the ToR.

Each stage of the project (FDR, FAT, delivery, assembly, SAT and training) is considered completed when both the Contractors and SOLARIS sign the appropriate protocol.

### 3.1.2 Design, source code and drawings

The Contractor of Part 1 is responsible for the dimensioning of the NAP-XPS system and its final operation, as well as fulfilling the guidelines from the Contractor of Part 2, regarding the design and manufacturing of the analysis chamber and differential pumping system for the operation of the Ambient Pressure Electron Energy Analyser. The detailed design will be based on a conceptual design.

The Contractor of Part 2 is responsible for the dimensioning of the Analyser module and its final operation, on condition that the Contractor of Part 1 meets the requirements for the analysis chamber in which the analyser will be mounted and the recommended set of pumps providing efficient differential pumping for the analyser. The Contractor of Part 2 is obliged to provide guidelines for the design of the analysis chamber within 1 month from the date of signing the Contract, including in the correspondence the representative of the Ordering Party indicated in the Contract.

The detailed design should include the NAP-XPS end station geometric layout based on calculations and simulations confirming the achievement of the specification requirements within the available space defined in the “Appendix –MECH3” and after consultations between Contractors. All drawings will be sent to SOLARIS as soon as they are made.

The system design must be submitted for the Ordering Party's approval in advance, allowing sufficient time to meet the deadline mentioned in § 4, section 1, point 1.1 of Appendix 2 to the ToR, while taking into account any comments from the Ordering Party on the design. The Ordering Party has two weeks to provide comments after receiving the complete design documentation (FDR) from the contractor. It is emphasized that SOLARIS verification will only concern the compliance / absence of collision with the existing / planned distribution of the beamline / the experimental hall space. The responsibility for achieving the required performance characteristics of the NAP-XPS system and ambient pressure electron energy analyser rests with the individual Contractors, who are responsible for their parts. The comments from the Ordering Party are in opinion-forming nature only. The Contractors are obliged to grant the Purchaser a license to use the technical documentation of the subject of the Contract, in particular for the purposes of usage, servicing and possible expansion of the delivered equipment, on the conditions specified in the Contract template. The technical documentation will be provided by the Contractors for SOLARIS use in a suitable editable digital format.

For Contractor of Part 1: The design will be presented in 2D and 3D mode and will include the location of Near Ambient Pressure XPS system components: chambers and vacuum components, HP cell module, frames, supports, infrastructure, etc.

For Contractor of Part 2: The design will be presented in 2D and 3D mode and will include the location of the Ambient Pressure Electron Energy Analyser components.

The Contractors will provide guidelines and detailed information (described below) which will be the base for the design and construction of the whole section infrastructure. All design guidelines for compressed air, water cooling, electricity, signal cables and IT infrastructure shall be delivered with the FDR meeting report.

The Contractors (for their part of the tender) shall present the schematic diagram of the electrical installation supplying the NAP-XPS end station components, e.g. electrical sockets for specific devices, alternating/constant voltage, active power. Racks power supply guidelines, grounding and equipotential bonding requirements, low voltage switchgears requirements shall also be provided. The electrical guidelines shall be prepared taking into account “Appendix EL1”. The devices described in “Appendix EL1” are not the subject of this order.

The Contractors (for their part of the tender) will provide guidelines for the low-voltage installation design – location and type of connection points at the devices, location of components of the NAP-XPS system (2D and 3D model).

The Contractors (for their part of the tender) will provide guidelines for the design of cooling water and compressed air installations (location of connection points at the devices, their types and required flows for proper devices operation and pressure drops, location of system components, 2D and 3D model). All parameters of the cooling water and compressed air system installed in SOLARIS is described in Appendix WAT-CW1 and WAT-CA1.

The Contractors are obliged to equip main analysis chamber and HP cell manipulator with reference sockets according to the standard provided by SOLARIS (see “Appendix ALIGN”). The location of reference points has to be agreed with SOLARIS at the design stage. With the delivery, the Contractors provide the documentation necessary to perform the positioning of the complete NAP-XPS end station, including the location of reference points with respect to the beam axis (fidualization documentation) determined in post-production measurements.

Positioning of the end station components and supports is to be made by the SOLARIS employees together with the Contractor of Part 1 based on the documentation delivered by the Contractor during the installation phase.

The Contractor of Part 1 shall provide a table of proposed motor and encoder models for each motorized axes, including range, repeatability, full step motion resolution and accuracy of each movement. Appendix CS1 contains details of the motion control standard at SOLARIS and all required data. All data required for the motorized axes configuration shall be delivered within 6 months from the design approval (FDR meeting) but not later than 6 weeks before delivery. All axes controlled with IcePAP controller should be prepared for connection of standard cabling defined in Appendix CS1. If necessary a dedicated patch panel should be provided.

The Contractor of Part 2 shall provide the API (Application Programming Interface) for the Windows 11 operating system (64-bit) to the program for controlling, collecting and storing data, simultaneously controlling the energy analyzer and detector. The Contractor shall provide an additional copy of the software for its integration with the Tango Controls 9.2+ system and for testing purposes.

The general requirements of the control system are described in Appendix CS0 and CS1.

All other requirements for the subject of the order are contained in the relevant Appendices and shall remain in force unless otherwise stated in this document.

## 3.2 Compliance with SOLARIS

The Near Ambient Pressure XPS system (Part 1 of the Tender) and Ambient Pressure Electron Analyser (Part 2 of the Tender) shall be delivered ready for connection with the systems and technologies used in SOLARIS. Wherever possible, standard elements (such as vacuum pumps and controllers, vacuum valves, gauges and their controllers, etc.) will comply with the standards described in appendices, unless otherwise stated in this document. Exceptions to this rule are only allowed when it is not possible to include standard components in the Contractor’s design. Such exceptions shall be clarified and discussed with the SOLARIS team. The proposed other technical solutions and manufacturing methods shall be submitted to SOLARIS in writing for approval. Approval, if granted, will be in writing.

## 3.3 Cooperation with SOLARIS

The Contractors will cooperate with SOLARIS with respect to the design of the individual system’s backbone infrastructure, i.e. power, water and compressed air supply installation, in order to maintain their consistency with the system design – also see the attached documents.

## 3.4 Quality control

Every three months the Contractors will prepare a short report describing the work progress, which will define the status of the design and manufacturing and, if they occur, encountered issues.

Upon request of SOLARIS, the Contractors will present original certificates of materials used, determining their composition, physical and chemical properties, mainly for materials which have to be used in the case of UHV subassemblies. Audit visits will be allowed during the components production.

## 3.5 Assembly Identification

The identification marking on the stainless-steel plate mounted on the vacuum chambers and analyser components, chemically etched or electrically engraved shall enable identification of each subassembly separately. The mark has to be visible and contain the subassembly drawing code.

## 3.6 Standard Tools

All standard tools necessary for the assembly of the system components shall be provided by the Contractors. Their list shall be agreed with SOLARIS in advance, before the system components are delivered.

## 3.7 Packaging and delivery

All subassemblies of the NAP-XPS end station shall be delivered to SOLARIS by the Contractors. All subassemblies of the NAP-XPS end station shall be protected against damage and pollution during shipment. All vacuum flanges should be secured with blinds or mechanical covers to protect the flange gaskets and to prevent contamination of the components. Where necessary, the system components shall be protected with suitable protective material. Adequate packaging and protection shall be provided. It shall be ensured that the system components being transported (with or without transport boxes) are stiff and that it is possible to lift them with a forklift or a crane without damaging them.

## 3.8 Local staff training and remote support

The personnel delegated by the Ordering Party shall be trained by the Contractors on proper operation, maintenance and safe servicing of the delivered equipment in accordance with the draft of the contract. The training shall be carried out during installation or not later than one month after the installation is completed. In addition, the Contractors will provide 24 hours per year of remote support (so called hot line) by telephone or via the Internet (e-mail/zoom, etc.) throughout the duration of the quality warranty.

## 3.9 Connecting to the backbone infrastructure

The installation of the NAP-XPS system is the responsibility of the Contractor of part 1, while the installation of the Analyser system is the responsibility of the Contractor of part 2. The Contractors will make all necessary electrical connections, connections for cooling water and compressed air between the system components and the backbone beamline infrastructure.

## 3.10 Elements not covered by the Contract

SOLARIS is obliged to provide:

a)	Radiological protection system for the Contractor’s staff that meets the requirements of local safety regulations as well as for the protection of radiation-sensitive equipment.

b)	Backbone infrastructure for the NAP-XPS end- station. In the contiguity of the end- station, SOLARIS will install all necessary IT network terminals, electrical switchgear for predefined sockets and electronic racks, cooling water and compressed air lines.

## 3.11 Factory Acceptance Tests (FAT)

The Contractor of Part 1 is obliged to perform the following tests at the factory (FAT):

- Leakproofness check and ultra-high vacuum check, including UHV chambers, valves, vacuum bellows, etc., according to values specified above and “Appendix VAC2”. A vacuum better than 5×10-10 mbar  shall be demonstrated as part of the FAT.

- The functionality of the high pressure cell as defined in section 2.1.3 shall be demonstrated during the FAT. The FAT shall include documentation of the maximum permissible pressure in the high pressure cell and the pressure in the analysis chamber at this pressure in the HP cell. Heating and cooling profiles shall be submitted as part of the FAT.

- The functionality of Cryogenic motorized 4-axes manipulator as defined in section 2.1.2 shall be demonstrated during FAT. Cooling to 10 K shall be demonstrated. Heating to 500°C in the case of a resistive heating system and 1200°C in the case of an e-beam heating system in UHV conditions shall be demonstrated. Heating and cooling profiles shall be submitted as part of the FAT.

- Residual Gas Analysis (RGA) test – according to “Appendix VAC2”.

- Functional tests for limit switches, electric motors, position encoders and all electromechanical components. It will be checked that all remotely operated devices work correctly. Compressed air servo-motors, motorized actuators, limit switches will be tested throughout the specified operating range.

- Accuracy tests, including all motorized and non-motorized movements, shall be performed considering displacement measurement, resolution and repeatability.

- All motions shall be tested both in air and under vacuum before delivery. Full stroke of movement, resolution and repeatability necessary for achieving full functionality shall be demonstrated. The supplier shall define the allowable parameter space for positions and movement of the motors.

- The water cooling system shall be tested with water at a pressure of 4 bar and maximum flow rates before delivery.

- The motorised and unmotorised axes motion tests will be carried out in regarding total travel movements and limit switches activation. Additionally for all motorized axes measurements of an accuracy, resolution and motion repeatability parameters shall be performed. The position of the limit switches will be measured in reference to the absolute encoder position, marker 0 (zero) of incremental encoder or reference switch and written to the report. The tests shall be conducted with mounted encoders and external measuring devices. All control lines (Home, Limit+, Limit-, Disable) should be tested with hardware compatible with control lines electrical interface shown in the Appendix CS1 Section 3.2.

The Purchaser shall be entitled to witness all tests that may be necessary under the terms of this specification (at its own expense), and shall be informed about the tests at least 3 weeks in advance. The test results must be documented and form an integral part of the required documentation.

The Contractor of Part 2 is obliged to perform the following tests at the factory (FAT):

- Leakproofness check and ultra-high vacuum check, including UHV chambers, valves, vacuum bellows, etc., according to values specified above and “Appendix VAC2”. The analyser module shall be installed in the UHV system in order to perform functional tests. A vacuum better than 5×10-10 mbar shall be demonstrated as part of the FAT.

- Residual Gas Analysis (RGA) test – according to “Appendix VAC2”.

- Functional tests of the ambient pressure electron energy analyser as defined in section 2.2 including all modes of operation shall be demonstrated during the FAT. The source of photons should be an external X-ray source. The scope of the tests will be defined during the FDR meeting.

## 3.12 Site Acceptance Tests (SAT)

For Contractor of Part 1, the final acceptance tests at the installation site performed after the assembly of all components of the NAP-XPS system will consist of:

- He leakproofness test and base pressure test – according to values specified above and  „Appendix VAC2”.

- Residual Gas Analysis (RGA) test – according to “Appendix VAC2”.

- Magnetic field test in the analysis chamber.

- Alignment test. The Contractor is required to check the setting/geometry of all components of the NAP-XPS system on the supports so that sample illumination with additional radiation sources and the transfer of sample holders are possible. The Contractor of pat 1 is obliged to cooperate with the Contractor of part 2 to check the analyser setting/geometry relative to the analysis chamber in order to ensure correct measurements performance.

- Electrical, hydraulic and pneumatic functionality test. Checking the mechanical/electrical functionality to defined limit values. Testing of the control system in various circumstances, such as: normal operation, power failure, simulated venting, simulated pump failure, simulated cooling system failure. Software and hardware testing for stepper motor controllers.

- Functional test of all devices with controllers/power supplies (for example sample heating, flood source, x-ray source, pumping system, bakeout system, etc.)

For Contractor of Part 2, the final acceptance tests at the installation site performed after the assembly of all components of the NAP-XPS system and Ambient pressure electron energy Analyser, will consist of:

- He leakproofness test and base pressure test – according to values specified above and „Appendix VAC2”.

- Residual Gas Analysis (RGA) test – according to “Appendix VAC2”.

- Alignment test. The Contractor of Part 2 is obliged to cooperate with the Contractor of part 1 to check the analyser setting/geometry relative to the analysis chamber in order to ensure correct measurements performance.

- Functionality tests of the analyser including all modes of operation by using the source of photons installed to the NAP-XPS system – standard X-ray source. The scope of the tests will be defined during the preparation of the FDR meeting.

- Functionality tests of the analyser including all modes of operation by using the photons delivered by the PHELIX beamline. The scope of the tests will be defined during the preparation of the FDR meeting. The tests will be conducted during the quality warranty period on a date specified by the Ordering Party after consultations with the Contractor.

# List of appendices

Appendix

- Appendix MECH1 – Mechanics

- Appendix MECH3 – Description of the end station area

- Appendix ALIGN – Guidelines in field of alignment

- Appendix WAT-CW1 – Cooling water basic standards

- Appendix WAT-CA1 –  Compressed air basic standards

- Appendix VAC1 – Guidelines for UHV Components at Solaris

- Appendix VAC2 – Technologies and materials for SOLARIS UHV devices

- Appendix CS0 – SOLARIS Control System Standards

- Appendix CS1 – Motion Control Standard

- Appendix EL1 – Electrical Requirements