Załącznik nr 8 do SWZ

Nr sprawy  Szp/ZP-015/2023

Opis Przedmiotu Zamówienia – urządzenia do sterylizacji

- Przegląd techniczny ma być przeprowadzony wg zaleceń producenta.

- Cena usługi przeglądu technicznego ma zawierać w sobie koszt usługi oraz dojazd serwisanta (wraz z materiałami eksploatacyjnymi tj. uszczelki, filtry, wężyki pomp detergentów).

- Zamawiający wymaga od wykonawcy:

- posiadania zaplecze techniczne (posiadanie dostępu do części zamiennych, części zużywalnych i materiałów eksploatacyjnych )

- dysponowania osobami posiadającymi określoną wiedzę oraz umiejętności przy wykonywaniu procedury przeglądów technicznych

- specjalistycznego sprzętu kontrolno-pomiarowego z ważnym certyfikatem legalizacji

- Wykonanie zalecanych przez producenta przeglądów technicznych, z częstotliwością wskazaną w postępowaniu, w siedzibie Zamawiającego, zgodnie z harmonogramem przeglądów.

- Umowa na 24 miesiące

- Zakres czynności serwisowych zgrzewarki w ramach przeglądu:

- sprawdzenie kabla zasilającego i wtyczki.

- sprawdzenie poprawności pracy wyświetlacza.

- demontaż obudowy oczyszczenie wnętrza urządzenia z kurzu i pozostałości materiałów zgrzewanych.

- sprawdzenie ślizgów teflonowych grzałek- w razie konieczności wymiana.

- sprawdzenie mocowań grzałek i odległości- w razie konieczności korekta ustawień.

- sprawdzenie kalibracji temperatury w razie konieczności kalibracja.

- sprawdzenie poprawności docisku, w razie konieczności korekta ustawień siły docisku.

- oczyszczenie powierzchni rolek dociskowych.

- sprawdzenie stanu rolek dociskowych w razie konieczności smarowanie lub wymiana.

- sprawdzenie stanu pasków napędowych, w razie konieczności wymiana.

- sprawdzenie połączeń kablowych dokręcenie łączników kabli.

- testy zgrzewu, ocena jakości i integralności zgrzewu, wykonanie testu zgrzewu paskowego i testu atramentowego (testy zapewnia wykonawca).

- testy siły rozrywu grzewu przeprowadzony specjalistycznym urządzeniem pomiarowym potwierdzone protokołem (przedstawić aktualny certyfikat wzorcowania urządzenia pomiarowego).

- kalibracja parametrów zgrzewu (temperatury, siły docisku i prędkości) przedstawić aktualne certyfikaty urządzeń pomiarowych, czynności kalibracyjne zakończone protokołem kalibracji.

- testy elektryczne

- sprawdzenie i ustawienia daty i godziny w zegarze systemowym.

- sprawdzenie taśmy drukarki- jeśli urządzenie posiada wbudowaną drukarkę.

- zerowanie komunikatu serwisowego-  jeśli wymagane.

- Zakres czynności serwisowych myjni Getinge w ramach przeglądu:

- sprawdzenie szczelności komory i orurowania.

- czyszczenie filtrów komory i filtrów wstępnych powietrza

- regulacji i sprawdzenie systemu drzwi

- sprawdzenie drukarki i czyszczenie głowicy.

- wymiana wężyków pomp perystaltycznych detergentów i przewodów detergentów.

- sprawdzenie ilości dozowania każdego z detergentów, kalibracja przepływomierzy (potwierdzone protokołem z kalibracji).

- sprawdzenie okablowania i dokręcenie połączeń elektrycznych.

- sprawdzenie elementów zabezpieczeń elektrycznych- test działania.

- kalibracja czujników temperatury komory i suszarki  (przedstawić aktualny certyfikat wzorcowania urządzenia pomiarowego) potwierdzone protokołem kalibracji na żądanie zamawiającego

- sprawdzenie poziomu wody w komorze , w razie potrzeby regulacja.

- potwierdzenie systemowych interwałów przeglądowych i zerowanie licznika przeglądów.

- sprawdzenie listy błędów z systemu sterowania.

- przeprowadzenie procesów testowych i kontrola parametrów pracy na wszystkich programach.

- sprawdzenie działania wyłączników bezpieczeństwa.

- oczyszczenie elementów przelotowych wózków wsadowych i ramion natryskowych, w tym wózka laparoskopowego, wymiana uszkodzonych elementów

- testy bezpieczeństwa medycznych urządzeń elektrycznych według normy PN-EN 62353 potwierdzone protokołami z pomiarów (przedstawić aktualny certyfikat wzorcowania urządzenia pomiarowego).

- na potwierdzenie prawidłowego przebiegu procesu: Wykonać pomiar rozkładu temperatur niezależnymi czujnikami (minimum 10 punktów pomiarowych temperatury wynik w postaci wydruków z pomiaru rozkładów temperatury dołączyć do dokumentacji z przeglądu.  Wymagane certyfikaty kalibracji instrumentów pomiarowych dołączyć do dokumentacji przeglądowej.

- Zakres czynności serwisowych sterylizatorów Getinge w ramach przeglądu:

- odmulanie wytwornicy pary.

- wymian filtra powietrza sterylnego komory.

- wymiana uszczelek drzwi komory

- sprawdzenie drukarki i czyszczenie głowicy.

- sprawdzenie działania zaworów bezpieczeństwa, test otwarcia przy zadanym ciśnieniu próbnym.

- sprawdzenie i regulacja ciśnienia roboczego generatora pary i systemów zabezpieczeń.

- oczyszczenie czujników poziomu wytwornicy pary, w razie konieczności wymiana uszczelnień.

- sprawdzenie elementów wysokoprądowych generatora pary.

- czyszczenie filtrów siatkowych.

- sprawdzenie i oczyszczenie odwadniaczy.

- sprawdzenie i regulacja wody chłodzącej do pompy próżniowej,

- wykonanie testu szczelności komory i elementów orurowania.

- sprawdzenie szczelności i drożności wymienników ciepła, odkamienianie wymienników.

- sprawdzenie działania wyłączników bezpieczeństwa.

- sprawdzenie działania zabezpieczenia drzwi., smarowanie elementów ruchomych, w razie konieczności regulacja,

- sprawdzenie działania sygnalizacji alarmów.

- Sprawdzenie poprawności wskazań czujników temperatury, ciśnienia i manometrów, w razie konieczności kalibracja zewnętrznymi przyrządami pomiarowymi, (przedstawić aktualny certyfikat wzorcowania urządzenia pomiarowego). Czynności kalibracyjne i sprawdzające potwierdzone protokołem kalibracji.

- sprawdzenie listy błędów z systemu sterowania.

- sprawdzenie szczelności grzałek, elektrod i osłon termicznych.

- sprawdzenie i regulacja poziomu wody w zbiornikach, w razie konieczności oczyszczenie zbiorników.

- przeprowadzenie procesów testowych i kontrola parametrów pracy na wszystkich programach sterylizacji.

- potwierdzenie systemowych interwałów przeglądowych i zerowanie licznika przeglądów.

- testy bezpieczeństwa medycznych urządzeń elektrycznych według normy PN-EN 62353 potwierdzone protokołami z pomiarów (przedstawić aktualny certyfikat wzorcowania urządzenia pomiarowego).

- wykonać  pomiar rozkładu temperatur i ciśnienia niezależnymi czujnikami (minimum 10 punktów pomiarowych temperatury i 1 ciśnienia).

- Po każdej czynności serwisowej wymagany wpis do Karty Technicznej Aparatu, zawierający datę wykonanej usługi, rodzaj wykonanej czynności, datę ważności przeglądu technicznego.

- Po każdej czynności serwisowej wymagane wystawienie raportu serwisowego i dostarczenie go do Sekcji Aparatury Medycznej.

- W przypadku naprawy urządzeń Zamawiającego będzie zlecał usługę serwisową osobnym zleceniem, fakturowanym oddzielnie.

- Okres gwarancji na wymienione części: 6 miesięcy.

- Sposób zleceń: e-mail lub telefonicznie

- Przeglądy wykonywane są w siedzibie Zamawiającego. W przypadku konieczności transportu aparatu do siedziby serwisu technicznego Wykonawcy, koszt transportu pokrywa Wykonawca.

Załącznik nr 8a do SWZ

Nr sprawy  Szp/ZP-015/2023

Opis Przedmiotu Zamówienia – fakoemulsyfikatory

            - Cena usługi przeglądu technicznego ma zawierać w sobie koszt usługi oraz dojazd serwisanta wraz z materiałami eksploatacyjnymi, naprawą urządzeń wraz z częściami zamiennymi.

            - Zamawiający wymaga od wykonawcy:

- posiadania zaplecze techniczne (posiadanie dostępu do części zamiennych, części zużywalnych i materiałów eksploatacyjnych )

- dysponowania osobami posiadającymi określoną wiedzę oraz umiejętności przy wykonywaniu procedury przeglądów technicznych

- dysponowania osobami posiadającymi imienne, aktualne certyfikaty na każde urządzenie, posiadającymi przeszkolenie w centrum szkoleniowym producenta (wymagane do okazania na życzenie zamawiającego)

- specjalistycznego sprzętu kontrolno-pomiarowego z ważnym certyfikatem legalizacji

            - Wykonanie zalecanych przez producenta przeglądów technicznych, z częstotliwością wskazaną w postępowaniu, w siedzibie Zamawiającego, zgodnie z harmonogramem przeglądów.

            - Zakres czynności serwisowych w ramach przeglądu: zgodnie z aktualną procedurą producenta (wykaz czynności tzw. check – lista stanowi załącznik nr 4 do umowy).

            - Po każdej czynności serwisowej wymagany wpis do Karty Technicznej Aparatu, zawierający datę wykonanej usługi, rodzaj wykonanej czynności, datę ważności przeglądu technicznego.

            - Po każdej czynności serwisowej wymagane wystawienie raportu serwisowego i dostarczenie go do Sekcji Aparatury Medycznej.

            - Czas reakcji po zgłoszeniu awarii (z wyłączeniem sobót oraz dni ustawowa wolnych od pracy) – do  48 godzin

Załącznik nr 4 do Umowy

Wykaz czynności serwisowych

Fakoemulsyfikator – Model: Constellation LXT

| CASE INFORMATION |  |
| --- | --- |
| Completed Using Instructions from V-QMS-0096939, Rev: |  |
| Case / Service Request #: |  |
| Customer: | Serial Number: |
| Software Version: |  |

NOTE: Throughout this data sheet, fill in the requested data and explain any unchecked boxes or empty fields in the Notes section at the end of the form. When filling out range fields, match the number of significant digits specified by the expected values. The decimal separator used throughout this document is the dot.

| CALIBRATED TEST EQUIPMENT |  |  |
| --- | --- | --- |
| Equipment | Calibration ID Number | Calibration Due Date |
| Scopemeter - N/A if not used |  |  |
| Pressure Meter - N/A if not used |  |  |
| Power Meter Interface - N/A if not used |  |  |
| Power Meter Detector Head, Thermopile - N/A if not used |  |  |
| Power Meter Detector Head, Photodiode - N/A if not used |  |  |
| Other ☐ N/A if not used |  |  |

|  | Select Type of Service: - SERVICE CALL ☐ PREVENTATIVE MAINTENANCE ☐ INSTALLATION |  |  |
| --- | --- | --- | --- |
| Select Applicable System: ☐ CONSTELLATION ☐ STANDALONE PUREPOINT LASER |  |  |  |
| Optional Features |  | Installed? |  |
| Illuminator: Phosphor or Xenon |  | - Phosphor ☐ Xenon | - N/A if Standalone Laser |
| Auxiliary Illuminator |  | - Yes ☐ No | - N/A if Standalone Laser |
| If Aux installed, do Illuminator types match? For example, either both Xenon OR both Phosphor? |  | - Yes ☐ No | - N/A if no Aux or Standalone Laser |
| Tray Arm |  | - Yes ☐ No | - N/A if Standalone Laser |
| Laser Embedded |  | - Yes ☐ No | - N/A if Standalone Laser |
| Laser Indirect Ophthalmoscope |  | - Yes ☐ No |  |
| Slit Lamp |  | - Yes ☐ No |  |

2.0 INSTALLATION - N/A – Standalone Laser or Not a Constellation Installation

- Plug in Back up Battery - Date and Time set for customer’s location - Entered customer name and location - Ensure Serial number matches system label - Select “Auxiliary Illuminator Installed” (☐ N/A - Auxiliary Illuminator NOT installed) - Select “HD Video Overlay”. ☐ N/A - HDVO NOT installed

3.0	PREVENTATIVE MAINTENANCE

- N/A - PM not performed, skip the PM Section.

### CONSTELLATION PM (☐ N/A, NOT A CONSTELLATION)

3.1.1.1.☐

3.1.1.2.☐

Inspect electrical and pneumatic handpiece connectors. (Replaced as needed)

Inspect Fan Filters. (Replaced as needed)

CONNECTORS AND FAN FILTERS

### TRAY ARM   (☐ N/A – No Tray Arm installed)

      - ☐	Tray Elbow, Tray/wrist and wrist/forearm joints are tight.

      - ☐	Inspect actuator pins, locking blades and hardware. Replace as needed.

      - ☐	Verify Actuator link set screw is tight

      - ☐	Verify Actuator travel. Adjust as necessary

    - FRONT PANEL MECHANICAL

      - ☐	Verify Display range of travel

### TOUCHSCREEN CALIBRATION

      - ☐	Touchscreen Calibration completed

### CPU BATTERY AND BACKUP BATTERY

      - Battery replaced	☐   Yes	☐   No

CONSTELLATION VIDEO/AUDIO & WIFI

      - ☐	Display output (☐ N/A, Display output not used)

      - ☐	Video Overlay Displayed (☐ N/A, Video Overlay not used)

      - ☐	DVD Recorder paused with footpedal (☐ N/A, DVD Recorder not used)

      - ☐	WiFI - Print End Case Report (☐ N/A, WiFi Printer not used)

3.2.1.☐

3.2.2.☐

3.2.3.☐

3.2.4.☐

3.2.5.☐

3.2.6.☐

Laser Goggles Inspected

Clean optics block components, if necessary Clean faces of fiber optics, if necessary Clean fans and TECs, if necessary

Check / set terminal fiber efficiencies Port 1 and 2, if necessary

Photomonitors calibration, if necessary

LASER PM (☐ N/A, NO LASER INSTALLED)

# 4.0	LASER FUNCTIONAL TEST (☐ N/A, NOT LASER)

    - ☐ Laser Safety Goggles Inspected (☐ N/A if already done on PM)

    - ☐ Emergency Stop functions

    - ☐ Filter message displayed

Filter not installed or passive filter used ☐Yes - 4.1.4 & 4.1.5 are N/A

      - No - Complete 4.1.4 & 4.1.5

    - ☐ Filter 1 lever – disengaged (☐ N/A - not equipped or passive filter used)

    - ☐ Filter 2 lever – disengaged (☐ N/A - not equipped or passive filter used)

    - ☐ Remote Interlock removed prevents firing

    - ☐ Footswitch removal displays message

    - ☐ LIO RFID Ports 1 & 2

    - ☐ LIO Illumination Maximum Brightness

    - ☐ Endoprobe RFID Ports 1 & 2

    - Exposure time	mS (9.5-10.5)

    - ENDOPROBE ENERGY VERIFICATION

Select measurement method: ☐ Energy OR ☐ Watts X Time

| Preset Power Value | Acceptable Power | Acceptable Energy | Endoprobe Port 1 Power | Endoprobe Port 1 Energy | Endoprobe Port 2 Power | Endoprobe Port 2 Energy |
| --- | --- | --- | --- | --- | --- | --- |
| 200 mW | 0.174 - 0.226 W | 1.64 - 2.36 mJ |  |  |  |  |
| 500 mW | 0.435 - 0.565 W | 4.1 – 5.9 mJ |  |  |  |  |
| 1.0 W | 0.870 - 1.13 W | 8.2 – 11.8 mJ |  |  |  |  |

    - ☐ Aiming Beam Maximum brightness

    - ☐ No connection prevents firing

    - STANDALONE CONSOLE CHECKS (☐ N/A - Not Standalone)

      - ☐ Inspect air filter. Clean or replace, if necessary

      - ☐ Interval knob

      - Record shot count

      - ☐ Check tethered capability - ☐ N/A if not tethered

| - LASER INDIRECT OPHTHALMOSCOPE (LIO) (☐ N/A - not equipped) - ☐ Illumination spot centered & proper size - ☐ Aiming Beam Alignment - ☐ Visual inspection of LIO Optical Fiber - LIO ENERGY VERIFICATION - Select measurement method: ☐ Energy OR ☐ Watts X Time Preset Acceptable Power Acceptable LIO LIO LIO LIO Power Energy Range Port 1 Port 1 Port 2 Port 2 Value Power Energy Power Energy 200 mW 0.174 - 0.226 W 1.64 - 2.36 mJ 500 mW 0.435 - 0.565 W 4.1 – 5.9 mJ 1.0 W 0.870 - 1.13 W 8.2 – 11.8 mJ - SLIT LAMP (☐ N/A, NOT EQUIPPED) |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 4.1.21.1 ☐ Proper operation of illumination and fixation bulbs |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.2 ☐ Free movement of joystick / chin rest / scope arm |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.3 ☐ Proper operation of arm and source locking screws |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.4 ☐ Free movement of magnification and filter knobs |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.5 ☐ Clean source / objective lens, mirror, and oculars |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.6 ☐ Converge / Focus / Center of slits at focal point |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.7 ☐ Converge / Focus of laser spot at focal point |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.8 ☐ Scope focus at all magnifications |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.9 ☐ Check movement of manipulator |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.10 ☐ Inspect Slit Lamp Optical Fiber. Replace if necessary. |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.11 ☐ Slit Lamp RFID Ports 1 & 2 |  |  |  |  |  |  |  |  |  |
|  | 4.1.21.12 ENERGY VERIFICATION - Select measurement method: ☐ Energy OR ☐ Watts X Time |  |  |  |  |  |  |  |  |  |
|  |  | Preset Power Value | Acceptable Power | Acceptable Energy Range | Slit Lamp Port 1 Power | Slit Lamp Port 1 Energy | Slit Lamp Port 2 Power | Slit Lamp Port 2 Energy |  |  |
|  |  | 200 mW | 0.174 - 0.226 W | 1.64 - 2.36 mJ |  |  |  |  |  |  |
|  |  | 500 mW | 0.435 - 0.565 W | 4.1 – 5.9 mJ |  |  |  |  |  |  |
|  |  | 1.0 W | 0.870 - 1.13 W | 8.2 – 11.8 mJ |  |  |  |  |  |  |

5.0	CONSTELLATION FUNCTIONAL TEST (☐ N/A, NOT A

CONSTELLATION)

### TRAY ARM (☐ N/A – No Tray Arm installed)

    - ☐ Tray arm range of travel.

### SYSTEM BOOTUP

    - ☐ System Bootup

### SOFTWARE

    - ☐ Check for and install software updates (If no update needed still check off this step)

    - Select System Software REL_ ☐ 4.01.14 ☐ 5.00.128 ☐ 5.40.26 ☐ Other - List and explain

5.4.2 Port 1

lumens (10 to 22)

5.4.3 Port 2

lumens (10 to 22)

5.4.4 ☐ RFID Ports 1 & 2

TABLETOP ILLUMINATION

  - AUXILIARY ILLUMINATOR (☐ N/A, NO AUXILIARY ILLUMINATOR INSTALLED)

5.5.2 Port 3

lumens (10 to 22)

5.5.3 Port 4

lumens (10 to 22)

5.5.4 ☐ RFID Ports 3 & 4

5.6.2 Diathermy Load

ohms (71.0 to 79.0)

5.6.5 Diathermy @ 100%

Vrms (24.5 to 30.0)

5.6.6 Diathermy Frequency

MHz (1.35 – 1.65)

DIATHERMY

  - AUTOSERT - ☐ N/A - NO AUTOSERT FUNCTION

5.7.6	☐	Autosert steps complete without error

### SMARTVIT MEASUREMENT SYSTEM TEST

5.8.6	☐	Verify date and time are correct.

5.8.2	☐	Ensure Serial number matches system label

5.8.46 ☐	Verify Datasheet displays “Passed”. Print to PDF and attach Smartvit Datasheet to service case.

### SYSTEM METRICS

| 5.9.2 ☐ Phosphor Illuminator, reset Lamp hours to 0. (☐ N/A – Xenon Illuminator) |  |  |
| --- | --- | --- |
| 5.9.3 Record Tabletop Xenon Lamp Hours |  | (☐ N/A – PHOSPHOR Illuminator) |
| 5.9.4 Record Auxiliary Xenon Lamp Hours Illuminator) |  | (☐ N/A - no Aux Illuminator or PHOSPHOR |
| 5.9.5 Record Laser Shots fired |  | (☐ N/A – Laser not installed) |

SD CARD READ/WRITE AND INSTRUMENT LOG

    - ☐ Review Instrument Log

    - ☐ Technicians Log copy without errors

    - ☐ Upload Log

### RING ILLUMINATION

    - SOFTWARE VERSIONS REL 5.00.XXX AND ABOVE. ☐ N/A - software is below 5.00.XXX.

      - Verify each ring illumination connector illuminates green, purple and amber.

### 5.11.4 SOFTWARE VERSIONS REL 4.01.XXX AND BELOW ☐ N/A if software is 5.00.XXX and above

- Verify each ring illumination connector illuminates green, purple and amber

### RFID

    - ☐	Verify each RFID connector illuminates green or amber.

### SCISSORS

5.13.4 Scissors	 	psi (48 to 52)

### FORCEPS

5.14.4 Forceps	 	psi (48 to 52)

### VFC INJECT

5.15.4 VFC Inject	 	psi (77 to 83)

VFC EXTRACT

    - ☐   VFC Extract @ Sea Level	 	mmHg (632 to 668)

OR

      - VFC Extract @ High Alt	 	mmHg (486 to 668)

  - AGF

5.17.5

5.17.7

5.17.10

5.17.11

SF6 Purge psi (29.5 to 30.5)

SF6 Fill psi (9.5 to 11.5)

C3F8 Purge psi (29.5 to 30.5)

C3F8 Fill psi (9.5 to 11.5)

### CASSETTE DETECTION AND ID

    - ☐	Cassette detection and ID

SUCTION

5.19.5 ☐   Suction @ Sea LEVEL	mmHg (632 to 668)

OR

- Suction @ High Alt	 	 mmHg (486 to 668)

  - FAX

5.20.2 FAX	mmHg (112.6 to 127.0)

### TOUCHSCREEN/TONE/VOICE

    - ☐Touchscreen/Tone/Voice

FOOTPEDAL

    - ☐ Footswitch Treadle and Switch Positions

    - ☐ Footswitch Vibration

5.22.4 ☐ Footswitch Proportional

  - REMOTE CONTROL ☐ N/A - NO REMOTE AVAILABLE

    - ☐	Remote Control functions

### BAR CODE SCANNER ☐ N/A - NO BAR CODE SCANNER AVAILABLE

5.24.2 ☐	Bar Code Scanner

### FLUIDIC PRIME

5.25.2 ☐	Successful Prime

PHACO

5.26.1 ☐ Phaco Power in each Phaco Panel Connectors

5.26.3 Record Phaco Handpiece Serial Number SN

  - FLOW TEST

5.27.2 ☐	Measured Flow matches Infusion Flow reading +/- 20%

Infusion

Backup Infusion

mmHg (112.6 to 127.0)

mmHg (18 to 42)

INFUSION / IRRIGATION

  - BACKUP BATTERY TEST

    - ☐ Power Recovery screen displayed and shuts down within 1 minute.

### MANUAL CASSETTE RELEASE

    - ☐ With Power off, verify manual cassette release functions

  - ELECTRICAL SAFETY TESTS ☐ N/A - ELECTRICAL SAFETY TESTS WERE NOT PERFORMED BY ALCON FIELD SERVICE.

    - Electrical Safety testing was performed and meets IEC/CEI 62353 standards or local standards.

| Equipment | Calibration ID Number | Calibration Due Date |
| --- | --- | --- |

Fakoemulsyfikator – Model: Centurion Vision System

Completed Using Instructions from V-QMS-0096601 Rev. No.

Case / Service Request #:

Customer:		Serial Number:

NOTE: Throughout this data sheet, fill in the requested data and explain any unchecked boxes or empty fields in the notes section at the end of the form. When filling out date fields, use the DD-MMM-YYYY format (e.g. 25-Jan-2020). When filling out range fields, match the number of significant digits specified by the expected values. The decimal separator used throughout this document is the dot.

# CALIBRATED TEST EQUIPMENT

| Equipment | Calibration ID Number | Calibration Due Date |
| --- | --- | --- |
| Scopemeter |  |  |
| Pressure Meter |  |  |

INSTALLATION

N/A - Not an installation, skip the remaining checks in this Installation section.

  - ☐ Remove cable tie from footswitch cable located in drawer

  - ☐ Check for and install software updates (if no upgrade needed still check off the step)

  - ☐ Date & Time set for the customer’s location

  - ☐ Entered local Technical Services phone number

  - ☐ Select country name based on installed location

  - ☐ Load language pack option (check off even if none are needed)

2.9.	☐ Check on whether or not the Active Sentry feature is to be enabled (even if not needed to be enabled check off the step)

  - ☐ Footswitch Network region set

  - ☐ Video Overlay Wi-Fi Network region set

  - ☐ Upload Network region set

PREVENTATIVE MAINTENANCE

N/A - PM not being performed, skip the remaining checks in this PM section.

Ω

≤0.5 Ohms

Power cord resistance

  - ☐ Display tilt and rotate operated correctly

  - ☐ Tray arm operated correctly

  - ☐ Work surface operated correctly

  - ☐ CPU Battery checked (replaced as needed)

  - ☐ Battery footswitch check (replaced as needed)

  - ☐ Console 12 volt batteries SOH is 50% or greater or they were replaced.

% Record SOH value

    - ☐ Verified exhaust air flowed from vents

    - ☐ Air filters are clean

  - ☐ Manufactured date and serial number are entered in system correctly

  - ☐ Cable wireless footswitch functioned correctly

  - ☐ Make certain all four casters are tight

FUNCTIONAL TEST

4.1.

4.2.

4.3.

4.4.

4.5.

# NO ADVISORY OR WARNING

- DATE AND TIME CORRECT

# TONES AND VOICE CONTROLS

- TASK LIGHT

# REMOTE CONTROL

- PEL LED PANELS TRACKED

# FOOTSWITCH TREADLE TRAVEL

- SIDE-SWITCHE

N/A ‐ NOT A CENTURION WIRELESS FOOTSWITCH, skip to 4.11.

  - ☐ 4 or more pedal_up sensors are functioning

  - Radio Strength for wireless footswitch communication at customer’s facility

| N/A ‐ Testing not performed at customer facility, skip to 4.8. Wireless Signal Strength | Expected Value | Displayed Value | Example |
| --- | --- | --- | --- |
| Diag. 3 | ≥ -40 |  | -30 is passing, -50 is failing |
| Diag. 4 | ≥ -40 |  | -30 is passing, -50 is failing |

4.10.5.

4.10.6.

# SOFTWARE:	REL		BLD

☐

☐

TOUCHSCREEN RESPONDED

# IV POLE FUNCTIONED

  - TRANSDUCER ACCURACY

| Input (mmHg) | Expected | Irr Pressure | Asp Pressure |
| --- | --- | --- | --- |
| +200 | +180 to +220 | + mmHg | + mmHg |
| -100 | -90 to -110 | ‐ mmHg | ‐ mmHg |
| -400 | -360 to -440 | ‐ mmHg | ‐ mmHg |

☐

CASSETTE PRIME

  - ASPIRATION FLOWRATE

|  | Result | Rate |
| --- | --- | --- |
| ☐ | 48 – 72 cc in 1 min. | @ 60 cc/min |
| ☐ | 20 – 30 cc in 1 min. | @ 25 cc/min |

# BAG PRESSURE SENSOR

|  | Result | Setting |
| --- | --- | --- |
| ☐ | 90 to 150 mmHg | 120 |
| ☐ | 140 to 200 mmHg | 170 |

  - PNEUMATIC PRESSURE TEST

| Bias | -0.6 to 1.2 PSI | PSI |
| --- | --- | --- |
| Peak Closed | ≥ 10.2 PSI | PSI |
| Peak Opened │absolute│ | ≥ │-10.1│ PSI | PSI |

# UNIVERSAL LOADBOX RESISTANCES

| Cautery Load | 71.0Ω to 79.0Ω | Ω |
| --- | --- | --- |
| Ozil Load | 2.37KΩ to 2.62KΩ | KΩ |

☐

ACTIVE SENTRY	N/A - Not an Active Sentry system, skip to 4.21.

| Pos. 1 from 1162 to 1176Ω | Ω |
| --- | --- |
| Pos. 2 from 1140 to 1154Ω | Ω |
| Pos. 3 from 1118 to 1132Ω | Ω |
| Pos. 4 from 1101 to 1115Ω | Ω |
| Pos. 5 from 1078 to 1092Ω | Ω |

# RECORD CAUTERY FREQUENCY AND VOLTAGES

| Frequency | 1.35 to 1.65 MHz | MHz |
| --- | --- | --- |
| 100% | 24.5 – 30.0 Vrms | Vrms |
| 50% | 17.3 – 21.2 Vrms | Vrms |
| 25% | 12.2 – 15.0 Vrms | Vrms |

  - ULTRASOUND OUTPUT TEST

| Top Long. Freq. | 41.6 – 42.2 KHz | KHz |
| --- | --- | --- |
| Top Long. Vrms | 9.9 – 13.3 Vrms | Vrms |
| Mid. Tors. Freq. | 31.7 – 32.2 KHz | KHz |
| Mid. Tors. Vrms | 4.9 – 6.6 Vrms | Vrms |
| Infiniti Long. Freq. | 33.5 – 37.5 KHz | KHz |
| Infiniti Long. Vrms | 9.9 – 13.3 Vrms | Vrms |
| Infiniti Tors. Freq. | 31.7 – 32.2 KHz | KHz |
| Infiniti Tors. Vrms | 4.9 – 6.6 Vrms | Vrms |

# ACTIVE SENTRY

- N/A - Not an Active Sentry System, skip to 4.24

| Top U/S Connector |  |
| --- | --- |
| - 268 to 297 | Position 1 |
| - 135 to 151 | Position 2 |
| - 2 to –2 | Position 3 |
| - –103 to –115 | Position 4 |
| - –243 to –269 | Position 5 |

| Middle U/S Connector |  |
| --- | --- |
| - 268 to 297 | Position 1 |
| - 135 to 151 | Position 2 |
| - 2 to –2 | Position 3 |
| - –103 to –115 | Position 4 |
| - –243 to –269 | Position 5 |

# ☐ IOL INJECTOR

  - ☐ EVENT LOG BACKED

# ☐ SPARE CENTURION WIRELESS FOOTSWITCH CABLE

  - ☐ BATTERY BACKUP

ELECTRICAL SAFETY TESTS

N/A - Electrical Safety Tests not performed by Alcon Field Service.

☐

| 5.2. Equipment | Calibration ID Number | Calibration Due Date |
| --- | --- | --- |

Electrical Safety testing was performed and meets IEC/CEI 62353 or equivalent local standards.

Zestaw do fakoemulsyfikacji – Model: Infiniti Vision System

Completed Using Instructions from V-QMS-0097184, Rev:

Case / Service Request #:

Customer:		Serial Number:

NOTE: Throughout this data sheet, fill in the requested data and explain any unchecked boxes or empty fields in the Notes section at the end of the form. When filling out date fields, use the DD-MMM-YYYY format (e.g. 25-Jan- 2020). When filling out range fields, match the number of significant digits specified by the expected values. The decimal separator used throughout this document is the dot.

# CALIBRATED TEST EQUIPMENT

| Equipment | Calibration ID Number | Calibration Due Date |
| --- | --- | --- |
| Scopemeter |  |  |
| Pressure Meter |  |  |

2.1.☐Event Viewer checked

PRE-TEST CHECKS

3.1.4.

Chassis resistance ≤ 0.5 ΩΩ

3.1.6.

Power cord resistance ≤ 0.5 ΩΩ

3.2.

☐

☐

☐

☐

☐

☐

☐

☐

☐

☐

☐

Display tilt and rotate operated correctly

Exhaust air flowed from vents

Air filter is clean

Casters

Fluidics Module Vent and Irr Spacers

Vent Solenoid shoulder screw

CPC Connector

CPU PCB battery

12 V Battery

Power Supplies

VideOverlay: typed characters present

PREVENTATIVE MAINTENANCE (☐ N/A, Preventative Maintenance not required. Continue to 4.)

FUNCTIONAL TEST

☐

☐

☐

No advisory or warning displayed

# Software:

REL_02.06

REL_03.02

☐

Date and time are correctly set in console

☐

Tones and voice controls worked correctly

4.5.1.☐

4.5.2.☐

Footswitch has full treadle travel motion

Side-switches are functioning properly

Footswitch:

  - Remote Control (☐ N/A, No remote present. Continue to 4.7.)

4.6.3.☐Remote control functions properly

☐

Touchscreen responded appropriately

# Transducer Accuracy:

| Input | Irrigation Specification | IPS Reading on Display | Aspiration Specification | APS Reading on Display |
| --- | --- | --- | --- | --- |
| +100 mmHg | +121 to +149 cmH2O | cmH2O | +90 to +110 mmHg | mmHg |
| -200 mmHg | -7 to +7 cmH2O | cmH2O | -180 to -220 mmHg | mmHg |
| -400 mmHg | -7 to +7 cmH2O | cmH2O | -360 to -440 mmHg | mmHg |

☐

Cassette passed priming process

# Aspiration Flowrate:

| Setting | Expected | Actual |
| --- | --- | --- |
| 60 cc/min | 48 – 72 cc | cc |
| 25 cc/min | 21 – 29 cc | cc |

  - Loadbox Verifications:

|  | Expected | Actual |
| --- | --- | --- |
| Cautery Load | 71.0 – 79.0 Ω | Ω |
| Ozil Load | 2.37 – 2.62 kΩ | kΩ |

# Pneumatics Test UltraVit 2500 CPM (☐ N/A, System does NOT have UltraVit. Continue to 4.11)

| Peak Closed | ≥ 10.2 PSI | PSI |
| --- | --- | --- |
| Peak Opened │absolute│ | ≥ │-10.1│ PSI | PSI |

Pressure out > 1.15 vpp

Pneumatics Test 800 CPM (☐ N/A, System DOES HAVE UltraVit, tested above. Continue to 4.12)

# Cautery Output:

|  | Expected | Actual |
| --- | --- | --- |
| Frequency | 1.43 – 1.67 MHz | MHz |
| 100% | 24.5 – 30.0 Vrms | Vrms |
| 50% | 17.3 – 21.2 Vrms | Vrms |
| 25% | 12.2 – 15.0 Vrms | Vrms |

  - Ultrasonic Output:

|  | Expected | Actual |
| --- | --- | --- |
| Top Freq. | 33.5 – 37.5 kHz | kHz |
| Top Vrms | 4.62 – 6.93 Vrms | Vrms |
| Bottom Freq. | 33.5 – 37.5 kHz | kHz |
| Bottom Vrms | 4.62 – 6.93 Vrms | Vrms |

4.14.1.9. ☐Checks complete, no error codes

IOL Injector / AutoSert Feature (☐ N/A, System does not have AutoSert. Continue to 4.15.)

☐

Event Log backed up

☐

Test shutdown

# ELECTRICAL SAFETY TESTS (☐ N/A, Electrical Safety Tests were not performed by Alcon Field Service.)

☐

5.1.		Electrical Safety testing was performed and meets IEC/CEI 62353 standards or equivalent local standards.

| Equipment | Calibration ID Number | Calibration Due Date |
| --- | --- | --- |