Proceedings **No. BZP.2710.5.2023.KWK**

**Appendix No. 3 to the ToR**

**Description of the subject matter of the order - Technical specification - Minimum requirements**

**Concerning the procedure entitled:** "Supply of an automatic bioaerosol detector with an atomiser and management and data visualisation software".

**Place of the contract delivery:** Department of Climatology and Atmospheric Protection, University of Wrocław, ul. Kosiby 8, 51-621 Wrocław.

|  |  |  |
| --- | --- | --- |
|  | **Detailed range of minimum technical parameters required by the Contracting Authority** | **Conformity of the features required with the features offered by the Contractor (To be filled in by the Contractor, by indicating YES or NO** as appropriate, e.g. for the answer YES: YES/~~NO~~ or YES/NO**, and in the blank space the Contractor shall clearly specify the parameters of the equipment/device/subassembly offered by him/her** |
| **General requirements** | * The device must use at least 3 pollen detection methods:   + **digital holography**   + **fluorescence spectroscopy (280 nm, 365 nm, 405 nm)**   + **polarisation measurements (405 nm)** | The device uses methods:  ………………………………………………………………………………………………………………………………………………………………………………………………………………….. |
|  | * The instrument must be equipped with an atomiser for expanding the database and teaching the recognition of new taxons | YES/NO\* |
|  | * The instrument must use a particle recognition algorithm based on artificial intelligence methods **(VGG 16)** | Recognition algorithm:  ……………………………………………………………………………………………………………………………………………………………… |
| **Time resolution of detection** | * At least 1 hour | Detection time resolution ...................h |
| **Minimum range of recognised particles** | * alnus, betula, carpinus, corylus, fagus, fraxinus, pinaceae, platanus, poaceae, populus, quercus, taxus, ulmus | Recognised particles:  ………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………. |
| **Functionality** | * Initial data, enabling monitoring to begin once the instrument is installed; * Access to the software code, allowing extension of the algorithms towards recognition of other taxa; * Software for management and visualisation of measurement data; | YES/NO\*  YES/NO\*  YES/NO\* |
| **Meteorological shield** | * Detector installed in a weatherproof enclosure. | YES/NO\* |